## Sheet1
Release_Number | S_No | Planning_Element | Question_No | Task_Owner_Contributing_Role | Impact_Analysis_Question | Response | Justification_or_Implementation_Details | Target_Completion_Workflow_Phase
GA2.1.2 | 1 | CHANGE IMPLEMENTATION BACKLOG | -- | Change Owner or Design Owner | Are there any design change implementation backlog? Identify any open (not COMPLETE) MR/ECR/ECO that must be RELEASED or IMPLEMENTED for this ECR to move forward. When multiple open MR/ECR are listed or when backlog impacts multiple sites, recommend documenting implementation strategy based on approved part disposition (e.g., sequence in which design change should be released or implemented) | NO | No open and not completed MR/ECR/ECO is pending prior to this Design change. | nan
GA2.1.2 | 2 | DESIGN HISTORY FILE | 1 | Engineering Representative | Does this design change require Product DHF (for the Products impacted by the proposed Design Change) to be updated? For example, Revise established product DHF Index if applicable. | YES | Update IPM DHF Index DOC2275693 for this Design Change release | PRIOR TO ECO RELEASE
GA2.1.2 | 3 | DESIGN INPUT | 2 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN INPUT design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact). There is no change to the Design Input | PRIOR TO ECO RELEASE
GA2.1.2 | 4 | DESIGN INPUT | 2.1 | Engineering Representative | Does this design change add or remove target markets and countries (e.g., impact on FDA, EU MDD/MDR, labeling and language requirements), requiring Design Input Requirements update? | NO | No change in Target Market and counties for this design change. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.2 | 5 | DESIGN INPUT | 2.2 | Engineering Representative Clinical Representative | Does this design change impact released Product User Needs and Intended Uses requirements included/referenced in DHF? | YES | Update IPM URS Document DOC2690834 to reflect the changes. | PRIOR TO ECO RELEASE
GA2.1.2 | 6 | DESIGN INPUT | 2.3 | Engineering Representative | Does this design change Impact released Product Performance and Functional design requirements (system/subsystem/component level requirements) included/referenced in DHF? | YES | Update IPM SRS DOC2690835 to reflect the changes. | PRIOR TO ECO RELEASE
GA2.1.2 | 7 | DESIGN INPUT | 2.6 | Engineering Representative | Does this design change impact applicable internal GEHC Global Requirements Specifications per GEHC_GQP_10.01.003_F001 GEHC Global Requirements Specifications Form (DOC2299885) established as part of product realization, for the product in scope? NOTE: The completed product GEHC Global Requirements Specifications Form is maintained in Product DHF Index. | NO | This change has no impact on GEHC Global Requirements Specifications. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.2 | 8 | DESIGN INPUT | 2.9 | Manufacturing Representative | Does this design change impact Manufacturing/manufacturing environment requirements (based on manufacturing capabilities and needs)? | NO | this change is done in a non impacted Cloud environement | nan
GA2.1.2 | 9 | DESIGN INPUT | 2.10 | Sourcing Representative | Does this design change impact Sourcing requirements (e.g., based on supplier capabilities and needs)? | NO | No new Souring requirements are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.2 | 10 | DESIGN INPUT | 2.11 | Engineering Representative Product Security Leader | Does this design change impact Product Privacy and Security requirements (e.g., Cyber threat model and controls, access controls, network and data security controls, product privacy and security manual instructions)? | NO | No update to Security threat models identified for this release as per scope items referenced in SW Change Set (Scope-Impact) worksheet. | nan
GA2.1.2 | 11 | DESIGN INPUT | 2.13 | Engineering Representative | If any other inputs are impacted, identify impacted work products and implementation task required. | NO | No other inputs are impacted for this release. | nan
GA2.1.2 | 12 | Failure Modes and Effects Analysis | 4.1 | Usability Representative | Does this design change impact released Task Analysis or Use Failure Mode and Effects Analysis (UFMEA) or requires creation of new UFMEA? | NO | No new UFMEA are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.2 | 13 | Failure Modes and Effects Analysis | 4.2 | Engineering Representative | Does this design change impact released Design FMEA (or other tools/methods, e.g., FTA, Engineering review and/or analysis) or requires developing new DFMEA records? | NO | No new FTA or other tools /methods are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.2 | 14 | Failure Modes and Effects Analysis | 4.3 | Manufacturing Representative | Does this design change impact released Process FMEA for manufacturing and/or used medical device processing or requires development of new PFMEA records? | NO | This change do not modify the Cloud process | nan
GA2.1.2 | 15 | REGULATORY ASSESSMENT | 5.6 | Manufacturing Representative | Does this design change impact procedures, product manufacturing processes, test specifications/reports, etc. that have regulatory impact? | NO | This change do not imply any procedure nor process modification. | nan
GA2.1.2 | 16 | REGULATORY ASSESSMENT | 5.11 | Regulatory Affairs Representative | Does this design change Introduce new commercial catalogs or offerings, requiring RA action? | NO | No new commercial offering or new catalogs are introduced for this design change. | nan
GA2.1.2 | 17 | REGULATORY ASSESSMENT | 5.16 | Engineering Representative | Other, identify impacted work products and implementation task required. | NO | No other work products and implementation task are required as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.2 | 18 | DESIGN OUTPUT | 6 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN OUTPUT design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and refer to VnV Plan & Reports for List of Design Output document updates. | PRIOR TO ECO RELEASE
GA2.1.2 | 19 | DESIGN OUTPUT | 6.6 | Engineering Representative Product Security Leader | Does this design change impact Cybersecurity Design Output deliverables: Impacts existing or requires new cybersecurity deliverables (e.g., DEPS Controls Baseline Catalog, Threat Model, SRA, PIA, MMSR, or Security Assessment Summary)? If yes, ensure PSR/PSL approvals are obtained as applicable | NO | There are no new Cybersecurity design changes identified as per scope deefined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.2 | 20 | DESIGN OUTPUT | 6.7 | Engineering Representative Product Security Leader | Does the design output updates in scope of this design change introduce new cybersecurity vulnerabilities that need to be assessed? Or are there any cybersecurity improvements that need to be implemented with this change? | NO | There are no new Security vulnerabilites identified for this design change. | nan
GA2.1.2 | 21 | DESIGN OUTPUT | 6.10 | Service Representative | Does this design change impact released product Service Manuals or documentation, e.g., service Tech Pub Labeling? For example, impact to a FRU/DRU part number or service instructions. Refer to Global QS Std WI for Technical Publications (DOC2139328) for further guidance | YES | IPM Service Manual will be updated prior to release of this design change. | PRIOR TO ECO RELEASE
GA2.1.2 | 22 | DESIGN OUTPUT | 6.11 | Installation Representative | Does this design change impact released product Installation manual/Pre-installation/de-installation/upgrade documentation (siting requirements, design center blocks and templates, customer drawings)? | YES | IPM User Manual will be updated prior to release of this design change. | PRIOR TO ECO RELEASE
GA2.1.2 | 23 | DESIGN OUTPUT | 6.13 | Engineering Representative | Does this design change impact released On Product Labeling/Operator/User Manuals (including addendums, datasheet, brochure or any other user facing instructions), e.g., User Tech Pub or Product Labeling/Documentation (Instructions for Use (IFU), translated addendums), product rating plate label, UDI label? | YES | IPM User Manual will be updated prior to release of this design change. | PRIOR TO ECO RELEASE
GA2.1.2 | 24 | DESIGN VERIFICATION | 7 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VERIFICATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | PRIOR TO ECO RELEASE
GA2.1.2 | 25 | DESIGN VERIFICATION | 7.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring verification to confirm that specified requirements are met, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Verification Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with verification results. NOTE: Update to Traceability of Design Input and lower-level requirements to associated Design Verification activities and results may be required. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | PRIOR TO ECO RELEASE
GA2.1.2 | 26 | DESIGN VERIFICATION | 7.3 | Engineering Representative | Does this design change require verification regression testing? | YES | Solution Verification Steering Guide will be prepared updated in later revision of this document during ECO. | PRIOR TO ECO RELEASE
GA2.1.2 | 27 | DESIGN VERIFICATION | 7.4 | Engineering Representative | Does this design change require technical analysis/specification analysis (e.g., scenarios where there is no impact to product design, however require engineering analysis to confirm this as part of justification) NOTE: Expert Evaluation and Analysis, including Technical Design Review, FMEA, FTA, and retrospective analysis of historical data applicable to the current design is one of the verification method and technique, per Verification GWI | YES | Scope limited to items referenced in SW Change Set (Scope & Impact). FMEA document will be updated prior ro release of Eco | PRIOR TO ECO RELEASE | 236
GA2.1.2 | 28 | DESIGN VERIFICATION | 7.6 | Engineering Representative | Does this design change impact Engineering test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for engineering development requiring re-verification/revision/re-validation/qualification? | NO | Scope limited to items referenced in SW Change Set (Scope & Impact). There is no additional tool, simulator or equipment update is required for this release. | nan | 352
GA2.1.2 | 29 | DESIGN VALIDATION | 8 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VALIDATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | PRIOR TO ECO RELEASE
GA2.1.2 | 30 | DESIGN VALIDATION | 8.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring validation to confirm that the product continues to meet user needs and intended uses, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Validation Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with validation results. NOTE: Update to Traceability of Design Input requirements to associated Design Validation activities and results may be required. NOTE: Consider all user groups when evaluating impact to validation, i.e., Service, Clinical, etc. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | PRIOR TO ECO RELEASE
GA2.1.2 | 31 | DESIGN VALIDATION | 8.2 | Engineering Representative | Does this design change require validation regression testing? | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | PRIOR TO ECO RELEASE
GA2.1.2 | 32 | SOFTWARE | 9 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to SOFTWARE design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Software Architecture Document will be updated Prior to release of this Design Change. | PRIOR TO ECO RELEASE
GA2.1.2 | 33 | SOFTWARE | 9.1 | Engineering Representative | Does this design change impact SOUP/third-party purchased software components? NOTE: If Yes, evaluate impact as per risk management and purchasing controls requirements. NOTE: If Yes, also verify to ensure that SOUP/third-party software functionality and no unintended side effects are introduced. | NO | Scope limited to items referenced in SW Change Set (Scope & Impact) and no NEW SOUP is required for this release. | nan
GA2.1.2 | 34 | SOFTWARE | 9.2 | Engineering Representative | Does the review of software (including SOUP) change determines whether: - Additional potential causes are introduced contributing to hazardous situations. - Additional risk mitigations are required. - The software change could interfere with existing risk mitigations. If Yes and there is an impact, identify relevant risk management activities based on impact of software design changes. | NO | No NEW SOUP is required for this release. | nan
GA2.1.2 | 35 | SOFTWARE | 9.3 | Engineering Representative | Does this design change require action per released product software maintenance plan (e.g., software architecture, SDD, software release, SOUP Index, software classification)? | YES | Software Architecture Document will be updated Prior to release of this Design Change. | PRIOR TO ECO RELEASE
GA2.1.2 | 36 | SOFTWARE | 9.4 | Engineering Representative | Does this design change include software upgrades, bug fixes, or obsolescence of SOUP items requiring software testing to ensure no unintended side effects are introduced? | NO | No change for this release. | nan
GA2.1.2 | 37 | SOFTWARE | 9.5 | Engineering Representative | Does this Design Change require released software validation test procedure to be executed to reestablish software validation? | NO | No change for this release. | nan
GA2.1.2 | 38 | SOFTWARE | 9.6 | Engineering Representative | Does this Design Change require updates to licensing servers/activation code servers based on software versioning? | NO | No change for this release. | nan
GA2.1.2 | 39 | USABILITY | 10 | Usability Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to USABILITY design controls element? If NO, document the rationale for no impact. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no change in Usability design control for this release. | nan
GA2.1.2 | 40 | USABILITY | 10.1 | Usability Representative | Does this design change impact Usability File Index (UFI) deliverables requiring update to UFI or Use Analysis Report (UAR)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.2 | 41 | USABILITY | 10.2 | Usability Representative Clinical Representative | Does this design change impact released Device Use Definition (including intended use, patient population, user profile, use environment, and primary operating functions and/or associated Use Scenarios (per UAR))? NOTE: Usability applies to interactions among the patient, User, device, and Use Environment. NOTE: Consider all user groups when evaluating impact to usability, i.e., Service, Clinical, etc. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.2 | 42 | USABILITY | 10.5 | Usability Representative | Does this design change impact Formative Evaluations and/or Summative Usability testing? If Yes, Identify action to create Summative Testing Plan and/or Formative Evaluation Plan as applicable. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.2 | 43 | USABILITY | 10.6 | Usability Representative | Does this design change impact User Interface of Unknown Provenance (UOUP) and is UOUP objective evidence needed (Device Use Definition in Use Analysis Report and documentation of Use Errors identified in Post-Market Data, including, but not limited to complaints, hazardous investigations with root cause of Use Error, near misses, adverse events, and recalls)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.2 | 44 | USABILITY | 10.7 | Usability Representative | Does this design change address Use Errors identified in complaints, adverse events, and/or recalls requiring a Known Use Problems evaluation? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.2 | 45 | DESIGN TRANSFER | 11.1 | Service Representative | Does this design change impact Service Finished design outputs? (e.g., Service manuals, service notes, asset creation in the service CRM, service checklist, replacement procedures, spare parts list, tools, processes, UDI Label Instructions, technical bulletin) Refer to QS_4004 Global Standard Work Instruction for Item Creation (DOC2473345) for further guidance on asset creation in service CRM. | YES | IPM Service Manual will be updated prior to release of ECO | PRIOR TO ECO RELEASE
GA2.1.2 | 46 | DESIGN TRANSFER | 11.2 | Service Representative | Does this Design Change require Installation & Service Process Validation? If yes, provide the installation & service process validation plan/procedure document number, and complete validation prior to ECO implementation. | YES | IPM Service Plan document number will be updated Prior to Release of this design change. | PRIOR TO ECO RELEASE
GA2.1.2 | 47 | DESIGN TRANSFER | 11.3 | Manufacturing Representative | Does this design change impact MAKE Parts? (i.e., Change to Manufacturing Process/Procedures/Work Instructions(WI)/OMS, SOP or require a New Process/Procedure, Process Validation, in-process label printing, testing, incoming, in-process, final acceptance, tooling, training, e.g., PPC or MVP), MTS, and PCL (Part Classification List) | YES | This change will impact SW PN and CAT# MAKE part. MR & ECO will be created. | PRIOR TO ECO RELEASE
GA2.1.2 | 48 | DESIGN TRANSFER | 11.8 | Engineering Representative | Does this design change impact established Product Monitoring design transfer strategy: Requires update to Product complaint handling process/system (consider impact to PSI codes and associated information, e.g., PSI Code, PSI Description, LEGAL MFG, MFG SITE, DESIGN OWNER)? OR...Requires activities to define content of reporting e.g. regulatory traceable device identification, Device History Record and Inventory Control, and Annual Report Identification and Set-Up? OR...Requires activities to define GEHC site interactions, e.g., update Site ASL with new GEHC Site, Quality Agreement, oversight and monitoring methodologies? | NO | No impact for this release | nan
GA2.1.2 | 49 | DESIGN TRANSFER | 11.9 | Manufacturing Representative | Does this design change affect eDHR or Paper DHR? | YES | a eDHR will be generated | PRIOR TO ECO RELEASE
GA2.1.2 | 50 | DESIGN TRANSFER | 11.13 | Supplier Quality Representative | Does PQP need to be updated for any other reason besides Essential controls? | NO | No essential control nor any PQP needed for reason beside Essential controls. | nan
GA2.1.2 | 51 | DESIGN TRANSFER | 11.16 | Manufacturing Representative | Does this design change impact Production Process Validation, Equipment qualification, Tooling or Fixture verification, Manufacturing Software, Training, Packaging and Shipping (e.g. labeling, shelf life, storage requirements). Impacts Material handling procedures, Environmental Requirements (Shelf Life, Temperature, Humidity Controls, etc.) New chemical introduced/SDS impacted, Manufacturing Process Capability | NO | No impact on production process and no associated material elements | nan
GA2.1.2 | 52 | TECHNICAL DESIGN REVIEW | 13 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to TECHNICAL DESIGN REVIEW design controls element? If NO, document the rationale for no impact. Refer to QS_4002 Global Standard WI for Technical Design Review (DOC0707915) for further guidance. | YES | TDR is planned, and all the required TDR plan, Calender , and TDRC-D-E document will be udated prior to release of ECO | PRIOR TO ECO RELEASE
GA2.1.2 | 53 | TECHNICAL DESIGN REVIEW | 13.1 | Engineering Representative | Does this proposed design change implementation require TDR to be planned and reviewed by engineering SME (e.g., Chief Engineer, LSD, PEs, Architects)? If yes identify TDR as part of TDR Calendar or implementation action plan and identify required SME based on topic. Examples may include Reliability TDR, Safety TDR, Usability TDR, Manufacturability TDR, Serviceability TDR, Software TDR, Compatibility TDR, Interoperability TDR. | YES | TDR is planned, and all the required TDR plan, Calender , and TDRC-D-E document will be udated prior to release of ECO | PRIOR TO ECO RELEASE
GA2.1.2 | 54 | TECHNICAL DESIGN REVIEW | 13.2 | Change Owner or Design Owner | Does this design change require TDR to be planned and executed per QS_3001 Global Standard WI for Suspended Mass Design (DOC1521734)? | YES | TDR is planned, and all the required TDR plan, Calender , and TDRC-D-E document will be udated prior to release of ECO | PRIOR TO ECO RELEASE
GA2.1.2 | 55 | TECHNICAL DESIGN REVIEW | 13.3 | Engineering Representative Product Security Leader | Does this design change require TDR to be planned and executed per GEHC New Product Cybersecurity Standard (DOC2102581)? | NO | No Security related TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.2 | 56 | TECHNICAL DESIGN REVIEW | 13.5 | Service Representative | Does this design change require TDR to be planned and executed per QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759)? If YES Service TDR Guidance and Checklist (DOC2444001) should be used. | NO | No Service TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.2 | 57 | Other RDOs | 15 | Engineering Representative | Does this design change impact another modality, system or subsystem managed by another change board If NO, document the rationale for no impact. If YES, indicate required tasks to complete. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no impact to other modalities , system, or sub-system managed by another change board. IPM Program manager will reach out to modality for announcement of this IPM Design change. | nan
GA2.1.2 | 58 | Platform Software or Hardware | 16 | Engineering Representative | Does this design change impact Platform software or hardware? If yes, identify design controls deliverables required to be completed prior to change implementation (e.g., platform integration and integration testing activities) If NO, document the rationale for no impact. | NO | IPM application is a Non-Medical Cloud based SaaS application, and it does impact any Platform software or hardware. Hence there is no impact. | nan
GA2.1.2 | 59 | OTHERS | nan | nan | nan | NO | nan | nan
GA2.1.3 | 60 | CHANGE IMPLEMENTATION BACKLOG | -- | Change Owner or Design Owner | Are there any design change implementation backlog? Identify any open (not COMPLETE) MR/ECR/ECO that must be RELEASED or IMPLEMENTED for this ECR to move forward. When multiple open MR/ECR are listed or when backlog impacts multiple sites, recommend documenting implementation strategy based on approved part disposition (e.g., sequence in which design change should be released or implemented) | NO | No open and not completed MR/ECR/ECO is pending prior to this Design change. | nan
GA2.1.3 | 61 | DESIGN HISTORY FILE | 1 | Engineering Representative | Does this design change require Product DHF (for the Products impacted by the proposed Design Change) to be updated? For example, Revise established product DHF Index if applicable. | YES | Update IPM DHF Index DOC2275693 for this Design Change release | nan
GA2.1.3 | 63 | DESIGN INPUT | 2.1 | Engineering Representative | Does this design change add or remove target markets and countries (e.g., impact on FDA, EU MDD/MDR, labeling and language requirements), requiring Design Input Requirements update? | NO | No change in Target Market and counties for this design change. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.3 | 64 | DESIGN INPUT | 2.2 | Engineering Representative Clinical Representative | Does this design change impact released Product User Needs and Intended Uses requirements included/referenced in DHF? | NO | No Change to User requirements for this DCP | nan
GA2.1.3 | 65 | DESIGN INPUT | 2.3 | Engineering Representative | Does this design change Impact released Product Performance and Functional design requirements (system/subsystem/component level requirements) included/referenced in DHF? | NO | No Change to System requirements for this DCP | nan
GA2.1.3 | 68 | DESIGN INPUT | 2.6 | Engineering Representative | Does this design change impact applicable internal GEHC Global Requirements Specifications per GEHC_GQP_10.01.003_F001 GEHC Global Requirements Specifications Form (DOC2299885) established as part of product realization, for the product in scope? NOTE: The completed product GEHC Global Requirements Specifications Form is maintained in Product DHF Index. | NO | This change has no impact on GEHC Global Requirements Specifications. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.3 | 71 | DESIGN INPUT | 2.9 | Manufacturing Representative | Does this design change impact Manufacturing/manufacturing environment requirements (based on manufacturing capabilities and needs)? | NO | this change is done in a non impacted Cloud environement | nan
GA2.1.3 | 72 | DESIGN INPUT | 2.10 | Sourcing Representative | Does this design change impact Sourcing requirements (e.g., based on supplier capabilities and needs)? | NO | No new Souring requirements are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.3 | 73 | DESIGN INPUT | 2.11 | Engineering Representative Product Security Leader | Does this design change impact Product Privacy and Security requirements (e.g., Cyber threat model and controls, access controls, network and data security controls, product privacy and security manual instructions)? | NO | No update to Security threat models identified for this release as per scope items referenced in SW Change Set (Scope-Impact) worksheet. | nan
GA2.1.3 | 75 | DESIGN INPUT | 2.13 | Engineering Representative | If any other inputs are impacted, identify impacted work products and implementation task required. | NO | No other inputs are impacted for this release. | nan
GA2.1.3 | 87 | Failure Modes and Effects Analysis | 4.1 | Usability Representative | Does this design change impact released Task Analysis or Use Failure Mode and Effects Analysis (UFMEA) or requires creation of new UFMEA? | NO | No new UFMEA are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.3 | 88 | Failure Modes and Effects Analysis | 4.2 | Engineering Representative | Does this design change impact released Design FMEA (or other tools/methods, e.g., FTA, Engineering review and/or analysis) or requires developing new DFMEA records? | NO | No new FTA or other tools /methods are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.3 | 89 | Failure Modes and Effects Analysis | 4.3 | Manufacturing Representative | Does this design change impact released Process FMEA for manufacturing and/or used medical device processing or requires development of new PFMEA records? | NO | This change do not modify the Cloud process | nan
GA2.1.3 | 98 | REGULATORY ASSESSMENT | 5.6 | Manufacturing Representative | Does this design change impact procedures, product manufacturing processes, test specifications/reports, etc. that have regulatory impact? | NO | This change do not imply any procedure nor process modification. | nan
GA2.1.3 | 103 | REGULATORY ASSESSMENT | 5.11 | Regulatory Affairs Representative | Does this design change Introduce new commercial catalogs or offerings, requiring RA action? | NO | No new commercial offering or new catalogs are introduced for this design change. | nan
GA2.1.3 | 108 | REGULATORY ASSESSMENT | 5.16 | Engineering Representative | Other, identify impacted work products and implementation task required. | NO | No other work products and implementation task are required as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.3 | 109 | DESIGN OUTPUT | 6 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN OUTPUT design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and refer to VnV Plan & Reports for List of Design Output document updates. | nan
GA2.1.3 | 110 | DESIGN OUTPUT | 6.1 | Marketing Representative | Does this design change Modify or releases a new commercial offering identified in PCM (Product Configuration Master), ACM (Accessories Configuration Master), ATO (Assemble To Order), PTO (Pick To Order) etc.? | NO | No change in the PCM , ACM, ATO and PTO for this design change release. | nan
GA2.1.3 | 115 | DESIGN OUTPUT | 6.6 | Engineering Representative Product Security Leader | Does this design change impact Cybersecurity Design Output deliverables: Impacts existing or requires new cybersecurity deliverables (e.g., DEPS Controls Baseline Catalog, Threat Model, SRA, PIA, MMSR, or Security Assessment Summary)? If yes, ensure PSR/PSL approvals are obtained as applicable | YES | Security Assessment Summary document will be updated for this design change release. | nan
GA2.1.3 | 116 | DESIGN OUTPUT | 6.7 | Engineering Representative Product Security Leader | Does the design output updates in scope of this design change introduce new cybersecurity vulnerabilities that need to be assessed? Or are there any cybersecurity improvements that need to be implemented with this change? | YES | SAS will be updated for this design change release. | nan
GA2.1.3 | 119 | DESIGN OUTPUT | 6.10 | Service Representative | Does this design change impact released product Service Manuals or documentation, e.g., service Tech Pub Labeling? For example, impact to a FRU/DRU part number or service instructions. Refer to Global QS Std WI for Technical Publications (DOC2139328) for further guidance | YES | IPM Service Manual will not be impacted but IPM Service Notes will be updated. | nan
GA2.1.3 | 120 | DESIGN OUTPUT | 6.11 | Installation Representative | Does this design change impact released product Installation manual/Pre-installation/de-installation/upgrade documentation (siting requirements, design center blocks and templates, customer drawings)? | NO | IPM Service Manual will not be impacted for this design change. | nan
GA2.1.3 | 122 | DESIGN OUTPUT | 6.13 | Engineering Representative | Does this design change impact released On Product Labeling/Operator/User Manuals (including addendums, datasheet, brochure or any other user facing instructions), e.g., User Tech Pub or Product Labeling/Documentation (Instructions for Use (IFU), translated addendums), product rating plate label, UDI label? | NO | IPM User Manual will NOT be updated for this design change. | nan
GA2.1.3 | 137 | DESIGN VERIFICATION | 7 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VERIFICATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.3 | 138 | DESIGN VERIFICATION | 7.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring verification to confirm that specified requirements are met, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Verification Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with verification results. NOTE: Update to Traceability of Design Input and lower-level requirements to associated Design Verification activities and results may be required. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.3 | 140 | DESIGN VERIFICATION | 7.3 | Engineering Representative | Does this design change require verification regression testing? | YES | Solution Verification Steering Guide will be prepared updated in later revision of this document during ECO. | nan
GA2.1.3 | 141 | DESIGN VERIFICATION | 7.4 | Engineering Representative | Does this design change require technical analysis/specification analysis (e.g., scenarios where there is no impact to product design, however require engineering analysis to confirm this as part of justification) NOTE: Expert Evaluation and Analysis, including Technical Design Review, FMEA, FTA, and retrospective analysis of historical data applicable to the current design is one of the verification method and technique, per Verification GWI | YES | Scope limited to items referenced in SW Change Set (Scope & Impact). FMEA document will be updated prior ro release of Eco | nan
GA2.1.3 | 143 | DESIGN VERIFICATION | 7.6 | Engineering Representative | Does this design change impact Engineering test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for engineering development requiring re-verification/revision/re-validation/qualification? | NO | Scope limited to items referenced in SW Change Set (Scope & Impact). There is no additional tool, simulator or equipment update is required for this release. | nan
GA2.1.3 | 144 | DESIGN VALIDATION | 8 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VALIDATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.3 | 145 | DESIGN VALIDATION | 8.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring validation to confirm that the product continues to meet user needs and intended uses, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Validation Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with validation results. NOTE: Update to Traceability of Design Input requirements to associated Design Validation activities and results may be required. NOTE: Consider all user groups when evaluating impact to validation, i.e., Service, Clinical, etc. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.3 | 146 | DESIGN VALIDATION | 8.2 | Engineering Representative | Does this design change require validation regression testing? | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.3 | 148 | SOFTWARE | 9.1 | Engineering Representative | Does this design change impact SOUP/third-party purchased software components? NOTE: If Yes, evaluate impact as per risk management and purchasing controls requirements. NOTE: If Yes, also verify to ensure that SOUP/third-party software functionality and no unintended side effects are introduced. | NO | Scope limited to items referenced in SW Change Set (Scope & Impact) and no NEW SOUP is required for this release. | nan
GA2.1.3 | 149 | SOFTWARE | 9.2 | Engineering Representative | Does the review of software (including SOUP) change determines whether: - Additional potential causes are introduced contributing to hazardous situations. - Additional risk mitigations are required. - The software change could interfere with existing risk mitigations. If Yes and there is an impact, identify relevant risk management activities based on impact of software design changes. | NO | No NEW SOUP is required for this release. | nan
GA2.1.3 | 150 | SOFTWARE | 9.3 | Engineering Representative | Does this design change require action per released product software maintenance plan (e.g., software architecture, SDD, software release, SOUP Index, software classification)? | NO | Software Architecture Document is not imapcted for this Design Change. | nan
GA2.1.3 | 151 | SOFTWARE | 9.4 | Engineering Representative | Does this design change include software upgrades, bug fixes, or obsolescence of SOUP items requiring software testing to ensure no unintended side effects are introduced? | NO | No change for this release. | nan
GA2.1.3 | 152 | SOFTWARE | 9.5 | Engineering Representative | Does this Design Change require released software validation test procedure to be executed to reestablish software validation? | NO | No change for this release. | nan
GA2.1.3 | 153 | SOFTWARE | 9.6 | Engineering Representative | Does this Design Change require updates to licensing servers/activation code servers based on software versioning? | NO | No change for this release. | nan
GA2.1.3 | 156 | USABILITY | 10 | Usability Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to USABILITY design controls element? If NO, document the rationale for no impact. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no change in Usability design control for this release. | nan
GA2.1.3 | 157 | USABILITY | 10.1 | Usability Representative | Does this design change impact Usability File Index (UFI) deliverables requiring update to UFI or Use Analysis Report (UAR)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.3 | 158 | USABILITY | 10.2 | Usability Representative Clinical Representative | Does this design change impact released Device Use Definition (including intended use, patient population, user profile, use environment, and primary operating functions and/or associated Use Scenarios (per UAR))? NOTE: Usability applies to interactions among the patient, User, device, and Use Environment. NOTE: Consider all user groups when evaluating impact to usability, i.e., Service, Clinical, etc. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.3 | 161 | USABILITY | 10.5 | Usability Representative | Does this design change impact Formative Evaluations and/or Summative Usability testing? If Yes, Identify action to create Summative Testing Plan and/or Formative Evaluation Plan as applicable. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.3 | 162 | USABILITY | 10.6 | Usability Representative | Does this design change impact User Interface of Unknown Provenance (UOUP) and is UOUP objective evidence needed (Device Use Definition in Use Analysis Report and documentation of Use Errors identified in Post-Market Data, including, but not limited to complaints, hazardous investigations with root cause of Use Error, near misses, adverse events, and recalls)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.3 | 163 | USABILITY | 10.7 | Usability Representative | Does this design change address Use Errors identified in complaints, adverse events, and/or recalls requiring a Known Use Problems evaluation? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.3 | 165 | DESIGN TRANSFER | 11.1 | Service Representative | Does this design change impact Service Finished design outputs? (e.g., Service manuals, service notes, asset creation in the service CRM, service checklist, replacement procedures, spare parts list, tools, processes, UDI Label Instructions, technical bulletin) Refer to QS_4004 Global Standard Work Instruction for Item Creation (DOC2473345) for further guidance on asset creation in service CRM. | YES | IPM Service Manual will NOT be updated Service notes will be updated for this desing change. | nan
GA2.1.3 | 166 | DESIGN TRANSFER | 11.2 | Service Representative | Does this Design Change require Installation & Service Process Validation? If yes, provide the installation & service process validation plan/procedure document number, and complete validation prior to ECO implementation. | NO | No IPM service plan is impacted for this design change release. | nan
GA2.1.3 | 167 | DESIGN TRANSFER | 11.3 | Manufacturing Representative | Does this design change impact MAKE Parts? (i.e., Change to Manufacturing Process/Procedures/Work Instructions(WI)/OMS, SOP or require a New Process/Procedure, Process Validation, in-process label printing, testing, incoming, in-process, final acceptance, tooling, training, e.g., PPC or MVP), MTS, and PCL (Part Classification List) | YES | This change will impact SW PN and CAT# MAKE part. MR & ECO will be created. | nan
GA2.1.3 | 172 | DESIGN TRANSFER | 11.8 | Engineering Representative | Does this design change impact established Product Monitoring design transfer strategy: Requires update to Product complaint handling process/system (consider impact to PSI codes and associated information, e.g., PSI Code, PSI Description, LEGAL MFG, MFG SITE, DESIGN OWNER)? OR...Requires activities to define content of reporting e.g. regulatory traceable device identification, Device History Record and Inventory Control, and Annual Report Identification and Set-Up? OR...Requires activities to define GEHC site interactions, e.g., update Site ASL with new GEHC Site, Quality Agreement, oversight and monitoring methodologies? | NO | No impact for this release | nan
GA2.1.3 | 173 | DESIGN TRANSFER | 11.9 | Manufacturing Representative | Does this design change affect eDHR or Paper DHR? | YES | A eDHR will be generated for this design change. | nan
GA2.1.3 | 177 | DESIGN TRANSFER | 11.13 | Supplier Quality Representative | Does PQP need to be updated for any other reason besides Essential controls? | NO | No essential control nor any PQP needed for reason beside Essential controls. | nan
GA2.1.3 | 180 | DESIGN TRANSFER | 11.16 | Manufacturing Representative | Does this design change impact Production Process Validation, Equipment qualification, Tooling or Fixture verification, Manufacturing Software, Training, Packaging and Shipping (e.g. labeling, shelf life, storage requirements). Impacts Material handling procedures, Environmental Requirements (Shelf Life, Temperature, Humidity Controls, etc.) New chemical introduced/SDS impacted, Manufacturing Process Capability | NO | No impact on production process and no associated material elements | nan
GA2.1.3 | 193 | TECHNICAL DESIGN REVIEW | 13.1 | Engineering Representative | Does this proposed design change implementation require TDR to be planned and reviewed by engineering SME (e.g., Chief Engineer, LSD, PEs, Architects)? If yes identify TDR as part of TDR Calendar or implementation action plan and identify required SME based on topic. Examples may include Reliability TDR, Safety TDR, Usability TDR, Manufacturability TDR, Serviceability TDR, Software TDR, Compatibility TDR, Interoperability TDR. | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.3 | 194 | TECHNICAL DESIGN REVIEW | 13.2 | Change Owner or Design Owner | Does this design change require TDR to be planned and executed per QS_3001 Global Standard WI for Suspended Mass Design (DOC1521734)? | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.3 | 195 | TECHNICAL DESIGN REVIEW | 13.3 | Engineering Representative Product Security Leader | Does this design change require TDR to be planned and executed per GEHC New Product Cybersecurity Standard (DOC2102581)? | NO | No Security related TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.3 | 197 | TECHNICAL DESIGN REVIEW | 13.5 | Service Representative | Does this design change require TDR to be planned and executed per QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759)? If YES Service TDR Guidance and Checklist (DOC2444001) should be used. | NO | No Service TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.3 | 204 | Other RDOs | 15 | Engineering Representative | Does this design change impact another modality, system or subsystem managed by another change board If NO, document the rationale for no impact. If YES, indicate required tasks to complete. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no impact to other modalities , system, or sub-system managed by another change board. IPM Program manager will reach out to modality for announcement of this IPM Design change. | nan
GA2.1.3 | 205 | Platform Software or Hardware | 16 | Engineering Representative | Does this design change impact Platform software or hardware? If yes, identify design controls deliverables required to be completed prior to change implementation (e.g., platform integration and integration testing activities) If NO, document the rationale for no impact. | NO | IPM application is a Non-Medical Cloud based SaaS application, and it does impact any Platform software or hardware. Hence there is no impact. | nan
GA2.1.3 | 206 | OTHERS | nan | nan | nan | NO | nan | nan
GA2.1.4 | 207 | CHANGE IMPLEMENTATION BACKLOG | -- | Change Owner or Design Owner | Are there any design change implementation backlog? Identify any open (not COMPLETE) MR/ECR/ECO that must be RELEASED or IMPLEMENTED for this ECR to move forward. When multiple open MR/ECR are listed or when backlog impacts multiple sites, recommend documenting implementation strategy based on approved part disposition (e.g., sequence in which design change should be released or implemented) | NO | No open and not completed MR/ECR/ECO is pending prior to this Design change. | nan
GA2.1.4 | 208 | DESIGN HISTORY FILE | 1 | Engineering Representative | Does this design change require Product DHF (for the Products impacted by the proposed Design Change) to be updated? For example, Revise established product DHF Index if applicable. | YES | Update IPM DHF Index DOC2275693 for this Design Change release | nan
GA2.1.4 | 210 | DESIGN INPUT | 2.1 | Engineering Representative | Does this design change add or remove target markets and countries (e.g., impact on FDA, EU MDD/MDR, labeling and language requirements), requiring Design Input Requirements update? | NO | No change in Target Market and counties for this design change. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.4 | 211 | DESIGN INPUT | 2.2 | Engineering Representative Clinical Representative | Does this design change impact released Product User Needs and Intended Uses requirements included/referenced in DHF? | NO | No Change to User requirements for this DCP | nan
GA2.1.4 | 212 | DESIGN INPUT | 2.3 | Engineering Representative | Does this design change Impact released Product Performance and Functional design requirements (system/subsystem/component level requirements) included/referenced in DHF? | NO | No Change to System requirements for this DCP | nan
GA2.1.4 | 215 | DESIGN INPUT | 2.6 | Engineering Representative | Does this design change impact applicable internal GEHC Global Requirements Specifications per GEHC_GQP_10.01.003_F001 GEHC Global Requirements Specifications Form (DOC2299885) established as part of product realization, for the product in scope? NOTE: The completed product GEHC Global Requirements Specifications Form is maintained in Product DHF Index. | NO | This change has no impact on GEHC Global Requirements Specifications. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.4 | 218 | DESIGN INPUT | 2.9 | Manufacturing Representative | Does this design change impact Manufacturing/manufacturing environment requirements (based on manufacturing capabilities and needs)? | NO | this change is done in a non impacted Cloud environement. | nan
GA2.1.4 | 219 | DESIGN INPUT | 2.10 | Sourcing Representative | Does this design change impact Sourcing requirements (e.g., based on supplier capabilities and needs)? | NO | No new Souring requirements are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.4 | 220 | DESIGN INPUT | 2.11 | Engineering Representative Product Security Leader | Does this design change impact Product Privacy and Security requirements (e.g., Cyber threat model and controls, access controls, network and data security controls, product privacy and security manual instructions)? | NO | No update to Security threat models identified for this release as per scope items referenced in SW Change Set (Scope-Impact) worksheet. | nan
GA2.1.4 | 222 | DESIGN INPUT | 2.13 | Engineering Representative | If any other inputs are impacted, identify impacted work products and implementation task required. | NO | No other inputs are impacted for this release. | nan
GA2.1.4 | 234 | Failure Modes and Effects Analysis | 4.1 | Usability Representative | Does this design change impact released Task Analysis or Use Failure Mode and Effects Analysis (UFMEA) or requires creation of new UFMEA? | NO | No new UFMEA are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.4 | 235 | Failure Modes and Effects Analysis | 4.2 | Engineering Representative | Does this design change impact released Design FMEA (or other tools/methods, e.g., FTA, Engineering review and/or analysis) or requires developing new DFMEA records? | NO | No new FTA or other tools /methods are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.4 | 236 | Failure Modes and Effects Analysis | 4.3 | Manufacturing Representative | Does this design change impact released Process FMEA for manufacturing and/or used medical device processing or requires development of new PFMEA records? | NO | This change do not modify the Cloud process | nan
GA2.1.4 | 237 | Failure Modes and Effects Analysis | 4.4 | Service Representative Installation Representative | Does this design change impact released Process FMEA for service, installation/de-installation/upgrade or requires development of new PFMEA records for service/installation/maintenance? | NO | Service is not impacted by this DCP, hence, no new PFMEA need to be introduce for this DCP release. | nan
GA2.1.4 | 245 | REGULATORY ASSESSMENT | 5.6 | Manufacturing Representative | Does this design change impact procedures, product manufacturing processes, test specifications/reports, etc. that have regulatory impact? | NO | This change do not imply any procedure nor process modification. | nan
GA2.1.4 | 250 | REGULATORY ASSESSMENT | 5.11 | Regulatory Affairs Representative | Does this design change Introduce new commercial catalogs or offerings, requiring RA action? | NO | No new commercial offering or new catalogs are introduced for this design change. | nan
GA2.1.4 | 255 | REGULATORY ASSESSMENT | 5.16 | Engineering Representative | Other, identify impacted work products and implementation task required. | NO | No other work products and implementation task are required as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.4 | 256 | DESIGN OUTPUT | 6 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN OUTPUT design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and refer to VnV Plan & Reports for List of Design Output document updates. | nan
GA2.1.4 | 257 | DESIGN OUTPUT | 6.1 | Marketing Representative | Does this design change Modify or releases a new commercial offering identified in PCM (Product Configuration Master), ACM (Accessories Configuration Master), ATO (Assemble To Order), PTO (Pick To Order) etc.? | NO | No change in the PCM , ACM, ATO and PTO for this design change release. | nan
GA2.1.4 | 262 | DESIGN OUTPUT | 6.6 | Engineering Representative Product Security Leader | Does this design change impact Cybersecurity Design Output deliverables: Impacts existing or requires new cybersecurity deliverables (e.g., DEPS Controls Baseline Catalog, Threat Model, SRA, PIA, MMSR, or Security Assessment Summary)? If yes, ensure PSR/PSL approvals are obtained as applicable | YES | Security Assessment Summary document will be updated for this design change release. | nan
GA2.1.4 | 263 | DESIGN OUTPUT | 6.7 | Engineering Representative Product Security Leader | Does the design output updates in scope of this design change introduce new cybersecurity vulnerabilities that need to be assessed? Or are there any cybersecurity improvements that need to be implemented with this change? | YES | SAST and DAST assessments will be performed for this design change release. | nan
GA2.1.4 | 266 | DESIGN OUTPUT | 6.10 | Service Representative | Does this design change impact released product Service Manuals or documentation, e.g., service Tech Pub Labeling? For example, impact to a FRU/DRU part number or service instructions. Refer to Global QS Std WI for Technical Publications (DOC2139328) for further guidance | YES | IPM Service Manual will be updated for this DCP release | nan
GA2.1.4 | 267 | DESIGN OUTPUT | 6.11 | Installation Representative | Does this design change impact released product Installation manual/Pre-installation/de-installation/upgrade documentation (siting requirements, design center blocks and templates, customer drawings)? | NO | IPM Service Manual will not be impacted for this design change. | nan
GA2.1.4 | 269 | DESIGN OUTPUT | 6.13 | Engineering Representative | Does this design change impact released On Product Labeling/Operator/User Manuals (including addendums, datasheet, brochure or any other user facing instructions), e.g., User Tech Pub or Product Labeling/Documentation (Instructions for Use (IFU), translated addendums), product rating plate label, UDI label? | NO | IPM User Manual will NOT be updated for this design change. | nan
GA2.1.4 | 284 | DESIGN VERIFICATION | 7 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VERIFICATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.4 | 285 | DESIGN VERIFICATION | 7.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring verification to confirm that specified requirements are met, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Verification Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with verification results. NOTE: Update to Traceability of Design Input and lower-level requirements to associated Design Verification activities and results may be required. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.4 | 287 | DESIGN VERIFICATION | 7.3 | Engineering Representative | Does this design change require verification regression testing? | YES | Solution Verification Steering Guide will be prepared updated in later revision of this document during ECO. | nan
GA2.1.4 | 288 | DESIGN VERIFICATION | 7.4 | Engineering Representative | Does this design change require technical analysis/specification analysis (e.g., scenarios where there is no impact to product design, however require engineering analysis to confirm this as part of justification) NOTE: Expert Evaluation and Analysis, including Technical Design Review, FMEA, FTA, and retrospective analysis of historical data applicable to the current design is one of the verification method and technique, per Verification GWI | YES | Scope limited to items referenced in SW Change Set (Scope & Impact). FMEA document will be updated prior ro release of Eco | nan
GA2.1.4 | 290 | DESIGN VERIFICATION | 7.6 | Engineering Representative | Does this design change impact Engineering test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for engineering development requiring re-verification/revision/re-validation/qualification? | NO | Scope limited to items referenced in SW Change Set (Scope & Impact). There is no additional tool, simulator or equipment update is required for this release. | nan
GA2.1.4 | 291 | DESIGN VALIDATION | 8 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VALIDATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.4 | 292 | DESIGN VALIDATION | 8.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring validation to confirm that the product continues to meet user needs and intended uses, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Validation Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with validation results. NOTE: Update to Traceability of Design Input requirements to associated Design Validation activities and results may be required. NOTE: Consider all user groups when evaluating impact to validation, i.e., Service, Clinical, etc. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.4 | 293 | DESIGN VALIDATION | 8.2 | Engineering Representative | Does this design change require validation regression testing? | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.4 | 295 | SOFTWARE | 9.1 | Engineering Representative | Does this design change impact SOUP/third-party purchased software components? NOTE: If Yes, evaluate impact as per risk management and purchasing controls requirements. NOTE: If Yes, also verify to ensure that SOUP/third-party software functionality and no unintended side effects are introduced. | NO | Scope limited to items referenced in SW Change Set (Scope & Impact) and no NEW SOUP is required for this release. | nan
GA2.1.4 | 296 | SOFTWARE | 9.2 | Engineering Representative | Does the review of software (including SOUP) change determines whether: - Additional potential causes are introduced contributing to hazardous situations. - Additional risk mitigations are required. - The software change could interfere with existing risk mitigations. If Yes and there is an impact, identify relevant risk management activities based on impact of software design changes. | NO | No NEW SOUP is required for this release. | nan
GA2.1.4 | 297 | SOFTWARE | 9.3 | Engineering Representative | Does this design change require action per released product software maintenance plan (e.g., software architecture, SDD, software release, SOUP Index, software classification)? | NO | Software Architecture Document is not imapcted for this Design Change. | nan
GA2.1.4 | 298 | SOFTWARE | 9.4 | Engineering Representative | Does this design change include software upgrades, bug fixes, or obsolescence of SOUP items requiring software testing to ensure no unintended side effects are introduced? | NO | No change for this release. | nan
GA2.1.4 | 299 | SOFTWARE | 9.5 | Engineering Representative | Does this Design Change require released software validation test procedure to be executed to reestablish software validation? | NO | No change for this release. | nan
GA2.1.4 | 300 | SOFTWARE | 9.6 | Engineering Representative | Does this Design Change require updates to licensing servers/activation code servers based on software versioning? | NO | No change for this release. | nan
GA2.1.4 | 303 | USABILITY | 10 | Usability Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to USABILITY design controls element? If NO, document the rationale for no impact. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no change in Usability design control for this release. | nan
GA2.1.4 | 304 | USABILITY | 10.1 | Usability Representative | Does this design change impact Usability File Index (UFI) deliverables requiring update to UFI or Use Analysis Report (UAR)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.4 | 305 | USABILITY | 10.2 | Usability Representative Clinical Representative | Does this design change impact released Device Use Definition (including intended use, patient population, user profile, use environment, and primary operating functions and/or associated Use Scenarios (per UAR))? NOTE: Usability applies to interactions among the patient, User, device, and Use Environment. NOTE: Consider all user groups when evaluating impact to usability, i.e., Service, Clinical, etc. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.4 | 308 | USABILITY | 10.5 | Usability Representative | Does this design change impact Formative Evaluations and/or Summative Usability testing? If Yes, Identify action to create Summative Testing Plan and/or Formative Evaluation Plan as applicable. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.4 | 309 | USABILITY | 10.6 | Usability Representative | Does this design change impact User Interface of Unknown Provenance (UOUP) and is UOUP objective evidence needed (Device Use Definition in Use Analysis Report and documentation of Use Errors identified in Post-Market Data, including, but not limited to complaints, hazardous investigations with root cause of Use Error, near misses, adverse events, and recalls)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.4 | 310 | USABILITY | 10.7 | Usability Representative | Does this design change address Use Errors identified in complaints, adverse events, and/or recalls requiring a Known Use Problems evaluation? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.4 | 312 | DESIGN TRANSFER | 11.1 | Service Representative | Does this design change impact Service Finished design outputs? (e.g., Service manuals, service notes, asset creation in the service CRM, service checklist, replacement procedures, spare parts list, tools, processes, UDI Label Instructions, technical bulletin) Refer to QS_4004 Global Standard Work Instruction for Item Creation (DOC2473345) for further guidance on asset creation in service CRM. | YES | IPM Service Manual will NOT be updated Service notes will be updated for this desing change. | nan
GA2.1.4 | 313 | DESIGN TRANSFER | 11.2 | Service Representative | Does this Design Change require Installation & Service Process Validation? If yes, provide the installation & service process validation plan/procedure document number, and complete validation prior to ECO implementation. | NO | No IPM service plan is impacted for this design change release. | nan
GA2.1.4 | 314 | DESIGN TRANSFER | 11.3 | Manufacturing Representative | Does this design change impact MAKE Parts? (i.e., Change to Manufacturing Process/Procedures/Work Instructions(WI)/OMS, SOP or require a New Process/Procedure, Process Validation, in-process label printing, testing, incoming, in-process, final acceptance, tooling, training, e.g., PPC or MVP), MTS, and PCL (Part Classification List) | YES | This change will impact SW PN and CAT# MAKE part. MR & ECO will be created. This change will also impact a Technical publications for which an ECR will be created | nan
GA2.1.4 | 319 | DESIGN TRANSFER | 11.8 | Engineering Representative | Does this design change impact established Product Monitoring design transfer strategy: Requires update to Product complaint handling process/system (consider impact to PSI codes and associated information, e.g., PSI Code, PSI Description, LEGAL MFG, MFG SITE, DESIGN OWNER)? OR...Requires activities to define content of reporting e.g. regulatory traceable device identification, Device History Record and Inventory Control, and Annual Report Identification and Set-Up? OR...Requires activities to define GEHC site interactions, e.g., update Site ASL with new GEHC Site, Quality Agreement, oversight and monitoring methodologies? | NO | No impact for this release | nan
GA2.1.4 | 320 | DESIGN TRANSFER | 11.9 | Manufacturing Representative | Does this design change affect eDHR or Paper DHR? | YES | A eDHR will be generated for this design change. | nan
GA2.1.4 | 324 | DESIGN TRANSFER | 11.13 | Supplier Quality Representative | Does PQP need to be updated for any other reason besides Essential controls? | NO | No essential control nor any PQP needed for reason beside Essential controls. | nan
GA2.1.4 | 327 | DESIGN TRANSFER | 11.16 | Manufacturing Representative | Does this design change impact Production Process Validation, Equipment qualification, Tooling or Fixture verification, Manufacturing Software, Training, Packaging and Shipping (e.g. labeling, shelf life, storage requirements). Impacts Material handling procedures, Environmental Requirements (Shelf Life, Temperature, Humidity Controls, etc.) New chemical introduced/SDS impacted, Manufacturing Process Capability | NO | No impact on production process and no associated material elements for this DCP release. | nan
GA2.1.4 | 340 | TECHNICAL DESIGN REVIEW | 13.1 | Engineering Representative | Does this proposed design change implementation require TDR to be planned and reviewed by engineering SME (e.g., Chief Engineer, LSD, PEs, Architects)? If yes identify TDR as part of TDR Calendar or implementation action plan and identify required SME based on topic. Examples may include Reliability TDR, Safety TDR, Usability TDR, Manufacturability TDR, Serviceability TDR, Software TDR, Compatibility TDR, Interoperability TDR. | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.4 | 341 | TECHNICAL DESIGN REVIEW | 13.2 | Change Owner or Design Owner | Does this design change require TDR to be planned and executed per QS_3001 Global Standard WI for Suspended Mass Design (DOC1521734)? | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.4 | 342 | TECHNICAL DESIGN REVIEW | 13.3 | Engineering Representative Product Security Leader | Does this design change require TDR to be planned and executed per GEHC New Product Cybersecurity Standard (DOC2102581)? | NO | No Security related TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.4 | 344 | TECHNICAL DESIGN REVIEW | 13.5 | Service Representative | Does this design change require TDR to be planned and executed per QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759)? If YES Service TDR Guidance and Checklist (DOC2444001) should be used. | NO | No Service TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.4 | 351 | Other RDOs | 15 | Engineering Representative | Does this design change impact another modality, system or subsystem managed by another change board If NO, document the rationale for no impact. If YES, indicate required tasks to complete. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no impact to other modalities , system, or sub-system managed by another change board. IPM Program manager will reach out to modality for announcement of this IPM Design change. | nan
GA2.1.4 | 352 | Platform Software or Hardware | 16 | Engineering Representative | Does this design change impact Platform software or hardware? If yes, identify design controls deliverables required to be completed prior to change implementation (e.g., platform integration and integration testing activities) If NO, document the rationale for no impact. | NO | IPM application is a Non-Medical Cloud based SaaS application, and it does impact any Platform software or hardware. Hence there is no impact. | nan
GA2.1.4 | 353 | OTHERS | nan | nan | nan | NO | nan | nan
GA2.1.5 | 354 | CHANGE IMPLEMENTATION BACKLOG | -- | Change Owner or Design Owner | Are there any design change implementation backlog? Identify any open (not COMPLETE) MR/ECR/ECO that must be RELEASED or IMPLEMENTED for this ECR to move forward. When multiple open MR/ECR are listed or when backlog impacts multiple sites, recommend documenting implementation strategy based on approved part disposition (e.g., sequence in which design change should be released or implemented) | NO | No open and not completed MR/ECR/ECO is pending prior to this Design change. | nan
GA2.1.5 | 355 | DESIGN HISTORY FILE | 1 | Engineering Representative | Does this design change require Product DHF (for the Products impacted by the proposed Design Change) to be updated? For example, Revise established product DHF Index if applicable. | YES | Update IPM DHF Index DOC2275693 for this Design Change release | nan
GA2.1.5 | 357 | DESIGN INPUT | 2.1 | Engineering Representative | Does this design change add or remove target markets and countries (e.g., impact on FDA, EU MDD/MDR, labeling and language requirements), requiring Design Input Requirements update? | NO | No change in Target Market and counties for this design change. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.5 | 358 | DESIGN INPUT | 2.2 | Engineering Representative Clinical Representative | Does this design change impact released Product User Needs and Intended Uses requirements included/referenced in DHF? | NO | No Change to User requirements for this DCP | nan
GA2.1.5 | 359 | DESIGN INPUT | 2.3 | Engineering Representative | Does this design change Impact released Product Performance and Functional design requirements (system/subsystem/component level requirements) included/referenced in DHF? | NO | No Change to System requirements for this DCP | nan
GA2.1.5 | 362 | DESIGN INPUT | 2.6 | Engineering Representative | Does this design change impact applicable internal GEHC Global Requirements Specifications per GEHC_GQP_10.01.003_F001 GEHC Global Requirements Specifications Form (DOC2299885) established as part of product realization, for the product in scope? NOTE: The completed product GEHC Global Requirements Specifications Form is maintained in Product DHF Index. | NO | This change has no impact on GEHC Global Requirements Specifications. This release Scope is limited to items referenced in SW Change Set (Scope-Impact) sheet only. | nan
GA2.1.5 | 363 | DESIGN INPUT | 2.7 | Service Representative Installation Representative | Does this design change impact released Installation or Service requirements (e.g., Installation, Maintenance and Repair, Tools, Diagnostic requirements, siting requirements)? If YES QS_1001_T001 Product Service Standards Assessment Summary and Justification Template (DOC2449236) should be reviewed/updated with the new requirements NOTE: Confirm if there are changes to the customer site requirements/customer room preparation, change to the PIM (Pre-Installation Manual) as it may in turn drive changes to GE Design Center content and ultimately customers who currently are reviewing drawings, budgeting for siting requirements and/or currently building their room. NOTE: Refer to QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759) for further guidance | NO | No changes to installation or service requirements. | nan
GA2.1.5 | 365 | DESIGN INPUT | 2.9 | Manufacturing Representative | Does this design change impact Manufacturing/manufacturing environment requirements (based on manufacturing capabilities and needs)? | NO | this change is done in a non impacted Cloud environement. | nan
GA2.1.5 | 366 | DESIGN INPUT | 2.10 | Sourcing Representative | Does this design change impact Sourcing requirements (e.g., based on supplier capabilities and needs)? | NO | No new Souring requirements are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.5 | 367 | DESIGN INPUT | 2.11 | Engineering Representative Product Security Leader | Does this design change impact Product Privacy and Security requirements (e.g., Cyber threat model and controls, access controls, network and data security controls, product privacy and security manual instructions)? | NO | No update to Security threat models identified for this release as per scope items referenced in SW Change Set (Scope-Impact) worksheet. | nan
GA2.1.5 | 369 | DESIGN INPUT | 2.13 | Engineering Representative | If any other inputs are impacted, identify impacted work products and implementation task required. | NO | No other inputs are impacted for this release. | nan
GA2.1.5 | 381 | Failure Modes and Effects Analysis | 4.1 | Usability Representative | Does this design change impact released Task Analysis or Use Failure Mode and Effects Analysis (UFMEA) or requires creation of new UFMEA? | NO | No new UFMEA are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.5 | 382 | Failure Modes and Effects Analysis | 4.2 | Engineering Representative | Does this design change impact released Design FMEA (or other tools/methods, e.g., FTA, Engineering review and/or analysis) or requires developing new DFMEA records? | NO | No new FTA or other tools /methods are identified for this release as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.5 | 383 | Failure Modes and Effects Analysis | 4.3 | Manufacturing Representative | Does this design change impact released Process FMEA for manufacturing and/or used medical device processing or requires development of new PFMEA records? | NO | This change do not modify the Cloud process | nan
GA2.1.5 | 384 | Failure Modes and Effects Analysis | 4.4 | Service Representative Installation Representative | Does this design change impact released Process FMEA for service, installation/de-installation/upgrade or requires development of new PFMEA records for service/installation/maintenance? | NO | Service is not impacted by this DCP, hence, no new PFMEA need to be introduce for this DCP release. | nan
GA2.1.5 | 389 | REGULATORY ASSESSMENT | 5.3 | Engineering Representative | Does this design change impact Labeling (e.g., updates to system labels, rating plates, regulatory/safety labels, instructions for use, indications for use, addition/deletion of contraindications, GTIN)? NOTE: Change in rating plate may have UDI implications (if it is a new medical device, ensure the medical device is structured at or below the rating plate level in the BOM). | YES | IPM User Manual will be updated as part of this DCP release. | nan
GA2.1.5 | 392 | REGULATORY ASSESSMENT | 5.6 | Manufacturing Representative | Does this design change impact procedures, product manufacturing processes, test specifications/reports, etc. that have regulatory impact? | NO | This change do not imply any procedure nor process modification. | nan
GA2.1.5 | 397 | REGULATORY ASSESSMENT | 5.11 | Regulatory Affairs Representative | Does this design change Introduce new commercial catalogs or offerings, requiring RA action? | NO | No new commercial offering or new catalogs are introduced for this design change. | nan
GA2.1.5 | 402 | REGULATORY ASSESSMENT | 5.16 | Engineering Representative | Other, identify impacted work products and implementation task required. | NO | No other work products and implementation task are required as per scope defined in SW Change Set (Scope-Impact) sheet. | nan
GA2.1.5 | 403 | DESIGN OUTPUT | 6 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN OUTPUT design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and refer to VnV Plan & Reports for List of Design Output document updates. | nan
GA2.1.5 | 404 | DESIGN OUTPUT | 6.1 | Marketing Representative | Does this design change Modify or releases a new commercial offering identified in PCM (Product Configuration Master), ACM (Accessories Configuration Master), ATO (Assemble To Order), PTO (Pick To Order) etc.? | NO | No change in the PCM , ACM, ATO and PTO for this design change release. | nan
GA2.1.5 | 409 | DESIGN OUTPUT | 6.6 | Engineering Representative Product Security Leader | Does this design change impact Cybersecurity Design Output deliverables: Impacts existing or requires new cybersecurity deliverables (e.g., DEPS Controls Baseline Catalog, Threat Model, SRA, PIA, MMSR, or Security Assessment Summary)? If yes, ensure PSR/PSL approvals are obtained as applicable | YES | Security Assessment Summary document will be updated for this design change release. | nan
GA2.1.5 | 410 | DESIGN OUTPUT | 6.7 | Engineering Representative Product Security Leader | Does the design output updates in scope of this design change introduce new cybersecurity vulnerabilities that need to be assessed? Or are there any cybersecurity improvements that need to be implemented with this change? | YES | SAST and DAST assessments will be performed for this design change release. | nan
GA2.1.5 | 413 | DESIGN OUTPUT | 6.10 | Service Representative | Does this design change impact released product Service Manuals or documentation, e.g., service Tech Pub Labeling? For example, impact to a FRU/DRU part number or service instructions. Refer to Global QS Std WI for Technical Publications (DOC2139328) for further guidance | YES | IPM Service Manual, Technical Overview & Customer IT, Service Release Note Requirments will be updated for this DCP release | nan
GA2.1.5 | 414 | DESIGN OUTPUT | 6.11 | Installation Representative | Does this design change impact released product Installation manual/Pre-installation/de-installation/upgrade documentation (siting requirements, design center blocks and templates, customer drawings)? | YES | IPM Service Manual will be impacted for this design change. | nan
GA2.1.5 | 416 | DESIGN OUTPUT | 6.13 | Engineering Representative | Does this design change impact released On Product Labeling/Operator/User Manuals (including addendums, datasheet, brochure or any other user facing instructions), e.g., User Tech Pub or Product Labeling/Documentation (Instructions for Use (IFU), translated addendums), product rating plate label, UDI label? | NO | IPM User Manual update is not required for this DCP with reference to the scope of this DCP. | nan
GA2.1.5 | 431 | DESIGN VERIFICATION | 7 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VERIFICATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.5 | 432 | DESIGN VERIFICATION | 7.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring verification to confirm that specified requirements are met, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Verification Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with verification results. NOTE: Update to Traceability of Design Input and lower-level requirements to associated Design Verification activities and results may be required. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Verification report will be updated post formal verification cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.5 | 434 | DESIGN VERIFICATION | 7.3 | Engineering Representative | Does this design change require verification regression testing? | YES | Solution Verification Steering Guide will be prepared updated in later revision of this document during ECO. | nan
GA2.1.5 | 435 | DESIGN VERIFICATION | 7.4 | Engineering Representative | Does this design change require technical analysis/specification analysis (e.g., scenarios where there is no impact to product design, however require engineering analysis to confirm this as part of justification) NOTE: Expert Evaluation and Analysis, including Technical Design Review, FMEA, FTA, and retrospective analysis of historical data applicable to the current design is one of the verification method and technique, per Verification GWI | YES | Scope limited to items referenced in SW Change Set (Scope & Impact). FMEA document will be updated prior ro release of Eco | nan
GA2.1.5 | 437 | DESIGN VERIFICATION | 7.6 | Engineering Representative | Does this design change impact Engineering test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for engineering development requiring re-verification/revision/re-validation/qualification? | NO | Scope limited to items referenced in SW Change Set (Scope & Impact). There is no additional tool, simulator or equipment update is required for this release. | nan
GA2.1.5 | 438 | DESIGN VALIDATION | 8 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VALIDATION design controls element? If NO, document the rationale for no impact. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.5 | 439 | DESIGN VALIDATION | 8.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring validation to confirm that the product continues to meet user needs and intended uses, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Validation Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with validation results. NOTE: Update to Traceability of Design Input requirements to associated Design Validation activities and results may be required. NOTE: Consider all user groups when evaluating impact to validation, i.e., Service, Clinical, etc. | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.5 | 440 | DESIGN VALIDATION | 8.2 | Engineering Representative | Does this design change require validation regression testing? | YES | Scope limited to items referenced in SW Change Set (Scope & Impact) and Validation report will be updated post formal Validation cycle. Refer VnV Plan & Report sheet for more details. | nan
GA2.1.5 | 442 | SOFTWARE | 9.1 | Engineering Representative | Does this design change impact SOUP/third-party purchased software components? NOTE: If Yes, evaluate impact as per risk management and purchasing controls requirements. NOTE: If Yes, also verify to ensure that SOUP/third-party software functionality and no unintended side effects are introduced. | NO | Scope limited to items referenced in SW Change Set (Scope & Impact) and no NEW SOUP is required for this release. | nan
GA2.1.5 | 443 | SOFTWARE | 9.2 | Engineering Representative | Does the review of software (including SOUP) change determines whether: - Additional potential causes are introduced contributing to hazardous situations. - Additional risk mitigations are required. - The software change could interfere with existing risk mitigations. If Yes and there is an impact, identify relevant risk management activities based on impact of software design changes. | NO | No NEW SOUP is required for this release. | nan
GA2.1.5 | 444 | SOFTWARE | 9.3 | Engineering Representative | Does this design change require action per released product software maintenance plan (e.g., software architecture, SDD, software release, SOUP Index, software classification)? | NO | Software Architecture Document is not imapcted for this Design Change. | nan
GA2.1.5 | 445 | SOFTWARE | 9.4 | Engineering Representative | Does this design change include software upgrades, bug fixes, or obsolescence of SOUP items requiring software testing to ensure no unintended side effects are introduced? | NO | No change for this release. | nan
GA2.1.5 | 446 | SOFTWARE | 9.5 | Engineering Representative | Does this Design Change require released software validation test procedure to be executed to reestablish software validation? | NO | No change for this release. | nan
GA2.1.5 | 447 | SOFTWARE | 9.6 | Engineering Representative | Does this Design Change require updates to licensing servers/activation code servers based on software versioning? | NO | No change for this release. | nan
GA2.1.5 | 450 | USABILITY | 10 | Usability Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to USABILITY design controls element? If NO, document the rationale for no impact. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no change in Usability design control for this release. | nan
GA2.1.5 | 451 | USABILITY | 10.1 | Usability Representative | Does this design change impact Usability File Index (UFI) deliverables requiring update to UFI or Use Analysis Report (UAR)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.5 | 452 | USABILITY | 10.2 | Usability Representative Clinical Representative | Does this design change impact released Device Use Definition (including intended use, patient population, user profile, use environment, and primary operating functions and/or associated Use Scenarios (per UAR))? NOTE: Usability applies to interactions among the patient, User, device, and Use Environment. NOTE: Consider all user groups when evaluating impact to usability, i.e., Service, Clinical, etc. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.5 | 455 | USABILITY | 10.5 | Usability Representative | Does this design change impact Formative Evaluations and/or Summative Usability testing? If Yes, Identify action to create Summative Testing Plan and/or Formative Evaluation Plan as applicable. | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.5 | 456 | USABILITY | 10.6 | Usability Representative | Does this design change impact User Interface of Unknown Provenance (UOUP) and is UOUP objective evidence needed (Device Use Definition in Use Analysis Report and documentation of Use Errors identified in Post-Market Data, including, but not limited to complaints, hazardous investigations with root cause of Use Error, near misses, adverse events, and recalls)? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.5 | 457 | USABILITY | 10.7 | Usability Representative | Does this design change address Use Errors identified in complaints, adverse events, and/or recalls requiring a Known Use Problems evaluation? | NO | No new User Feature added hence no change in Usablity applicable for this release. | nan
GA2.1.5 | 459 | DESIGN TRANSFER | 11.1 | Service Representative | Does this design change impact Service Finished design outputs? (e.g., Service manuals, service notes, asset creation in the service CRM, service checklist, replacement procedures, spare parts list, tools, processes, UDI Label Instructions, technical bulletin) Refer to QS_4004 Global Standard Work Instruction for Item Creation (DOC2473345) for further guidance on asset creation in service CRM. | YES | IPM Service Manual, Technical Overview & Customer IT, Service Release Note Requirments will be updated for this DCP release | nan
GA2.1.5 | 460 | DESIGN TRANSFER | 11.2 | Service Representative | Does this Design Change require Installation & Service Process Validation? If yes, provide the installation & service process validation plan/procedure document number, and complete validation prior to ECO implementation. | YES | Yes Service Validation will be occuring. | nan
GA2.1.5 | 461 | DESIGN TRANSFER | 11.3 | Manufacturing Representative | Does this design change impact MAKE Parts? (i.e., Change to Manufacturing Process/Procedures/Work Instructions(WI)/OMS, SOP or require a New Process/Procedure, Process Validation, in-process label printing, testing, incoming, in-process, final acceptance, tooling, training, e.g., PPC or MVP), MTS, and PCL (Part Classification List) | YES | This change will impact SW PN and CAT# MAKE part. MR & ECO will be created. | nan
GA2.1.5 | 466 | DESIGN TRANSFER | 11.8 | Engineering Representative | Does this design change impact established Product Monitoring design transfer strategy: Requires update to Product complaint handling process/system (consider impact to PSI codes and associated information, e.g., PSI Code, PSI Description, LEGAL MFG, MFG SITE, DESIGN OWNER)? OR...Requires activities to define content of reporting e.g. regulatory traceable device identification, Device History Record and Inventory Control, and Annual Report Identification and Set-Up? OR...Requires activities to define GEHC site interactions, e.g., update Site ASL with new GEHC Site, Quality Agreement, oversight and monitoring methodologies? | NO | No impact for this release | nan
GA2.1.5 | 467 | DESIGN TRANSFER | 11.9 | Manufacturing Representative | Does this design change affect eDHR or Paper DHR? | YES | A eDHR will be generated for this design change. | nan
GA2.1.5 | 471 | DESIGN TRANSFER | 11.13 | Supplier Quality Representative | Does PQP need to be updated for any other reason besides Essential controls? | NO | No essential control nor any PQP needed for reason beside Essential controls. | nan
GA2.1.5 | 474 | DESIGN TRANSFER | 11.16 | Manufacturing Representative | Does this design change impact Production Process Validation, Equipment qualification, Tooling or Fixture verification, Manufacturing Software, Training, Packaging and Shipping (e.g. labeling, shelf life, storage requirements). Impacts Material handling procedures, Environmental Requirements (Shelf Life, Temperature, Humidity Controls, etc.) New chemical introduced/SDS impacted, Manufacturing Process Capability | NO | No impact on production process and no associated material elements for this DCP release. | nan
GA2.1.5 | 481 | DESIGN TRANSFER | 11.23 | Service Representative | Does this design change impact Proprietary Classification of Hardware/Software Tools, Documentation, and Files documented in Product Proprietary Classification record per QS_10001 Global Standard WI for Proprietary Classification (DOC1410674)? | NO | No changes to the Product Proprietary Classification Record | nan
GA2.1.5 | 487 | TECHNICAL DESIGN REVIEW | 13.1 | Engineering Representative | Does this proposed design change implementation require TDR to be planned and reviewed by engineering SME (e.g., Chief Engineer, LSD, PEs, Architects)? If yes identify TDR as part of TDR Calendar or implementation action plan and identify required SME based on topic. Examples may include Reliability TDR, Safety TDR, Usability TDR, Manufacturability TDR, Serviceability TDR, Software TDR, Compatibility TDR, Interoperability TDR. | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.5 | 488 | TECHNICAL DESIGN REVIEW | 13.2 | Change Owner or Design Owner | Does this design change require TDR to be planned and executed per QS_3001 Global Standard WI for Suspended Mass Design (DOC1521734)? | NO | No new feature is planned for this design change, hence no change and impact to TDR plan. | nan
GA2.1.5 | 489 | TECHNICAL DESIGN REVIEW | 13.3 | Engineering Representative Product Security Leader | Does this design change require TDR to be planned and executed per GEHC New Product Cybersecurity Standard (DOC2102581)? | NO | No Security related TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.5 | 491 | TECHNICAL DESIGN REVIEW | 13.5 | Service Representative | Does this design change require TDR to be planned and executed per QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759)? If YES Service TDR Guidance and Checklist (DOC2444001) should be used. | NO | No Service TDR is required as per the scope defined in SW Change Set (Scope & Impact) sheet. | nan
GA2.1.5 | 498 | Other RDOs | 15 | Engineering Representative | Does this design change impact another modality, system or subsystem managed by another change board If NO, document the rationale for no impact. If YES, indicate required tasks to complete. | NO | Scope of this Design Change Release is limited to items as listed in SW Change Set (Scope & Impact) and there is no impact to other modalities , system, or sub-system managed by another change board. IPM Program manager will reach out to modality for announcement of this IPM Design change. | nan
GA2.1.5 | 499 | Platform Software or Hardware | 16 | Engineering Representative | Does this design change impact Platform software or hardware? If yes, identify design controls deliverables required to be completed prior to change implementation (e.g., platform integration and integration testing activities) If NO, document the rationale for no impact. | NO | IPM application is a Non-Medical Cloud based SaaS application, and it does impact any Platform software or hardware. Hence there is no impact. | nan
GA2.1.5 | 500 | OTHERS | nan | nan | nan | NO | nan | nan