## Sheet1
S_No | Planning_Element | Question_No | Task_Owner_Contributing_Role | Impact_Analysis_Question
1 | CHANGE IMPLEMENTATION BACKLOG | -- | Change Owner or Design Owner | Are there any design change implementation backlog? Identify any open (not COMPLETE) MR/ECR/ECO that must be RELEASED or IMPLEMENTED for this ECR to move forward. When multiple open MR/ECR are listed or when backlog impacts multiple sites, recommend documenting implementation strategy based on approved part disposition (e.g., sequence in which design change should be released or implemented)
2 | DESIGN HISTORY FILE | 1 | Engineering Representative | Does this design change require Product DHF (for the Products impacted by the proposed Design Change) to be updated? For example, Revise established product DHF Index if applicable.
3 | DESIGN INPUT | 2 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to "DESIGN INPUT" design controls element? If NO, document the rationale for no impact.
4 | DESIGN INPUT | 2.1 | Engineering Representative | Does this design change add or remove target markets and countries (e.g., impact on FDA, EU MDD/MDR, labeling and language requirements), requiring Design Input Requirements update?
5 | DESIGN INPUT | 2.2 | Engineering Representative Clinical Representative | Does this design change impact released Product User Needs and Intended Uses requirements included/referenced in DHF?
6 | DESIGN INPUT | 2.3 | Engineering Representative | Does this design change Impact released Product Performance and Functional design requirements (system/subsystem/component level requirements) included/referenced in DHF?
7 | DESIGN INPUT | 2.4 | Engineering Representative | Does this design change Impact product compatibility requirements (addition or removal of compatible devices or accessories)?
8 | DESIGN INPUT | 2.5 | Engineering Representative | Does this design change impact Product Standards design input requirements that apply to features, technology for the identified target markets and countries (e.g., IEC60601-1, ISO 10933, WEEE, RoHS, Batteries, Telecommunication, Wi-Fi, Bluetooth)?
9 | DESIGN INPUT | 2.6 | Engineering Representative | Does this design change impact applicable internal GEHC Global Requirements Specifications per GEHC_GQP_10.01.003_F001 GEHC Global Requirements Specifications Form (DOC2299885) established as part of product realization, for the product in scope? NOTE: The completed product GEHC Global Requirements Specifications Form is maintained in Product DHF Index.
10 | DESIGN INPUT | 2.7 | Service Representative Installation Representative | Does this design change impact released Installation or Service requirements (e.g., Installation, Maintenance and Repair, Tools, Diagnostic requirements, siting requirements)? If YES QS_1001_T001 Product Service Standards Assessment Summary and Justification Template (DOC2449236) should be reviewed/updated with the new requirements NOTE: Confirm if there are changes to the customer site requirements/customer room preparation, change to the PIM (Pre-Installation Manual) as it may in turn drive changes to GE Design Center content and ultimately customers who currently are reviewing drawings, budgeting for siting requirements and/or currently building their room. NOTE: Refer to QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759) for further guidance
11 | DESIGN INPUT | 2.8 | Engineering Representative | Does this design change impact Packaging requirements? (e.g., size, weight, storage, shelf life, packaging material, special handling)?
12 | DESIGN INPUT | 2.9 | Manufacturing Representative | Does this design change impact Manufacturing/manufacturing environment requirements (based on manufacturing capabilities and needs)?
13 | DESIGN INPUT | 2.10 | Sourcing Representative | Does this design change impact Sourcing requirements (e.g., based on supplier capabilities and needs)?
14 | DESIGN INPUT | 2.11 | Engineering Representative Product Security Leader | Does this design change impact Product Privacy and Security requirements (e.g., Cyber threat model and controls, access controls, network and data security controls, product privacy and security manual instructions)?
15 | DESIGN INPUT | 2.12 | Engineering Representative | Does this design change impact electronic interface interoperability requirements (e.g., DICOM, IHE, HL7 conformance)? NOTE: Refer to QS_5001 Global Standard WI for Conformity with Recognized Medical Interoperability Standards (DOC2432296) for further guidance
16 | DESIGN INPUT | 2.13 | Engineering Representative | If any other inputs are impacted, identify impacted work products and implementation task required.
17 | RISK MANAGEMENT | 3 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to RISK MANAGEMENT design controls element? If NO, document the rationale for no impact.
18 | RISK MANAGEMENT | 3.1 | Engineering Representative | Does this design change impact existing risk mitigations, or identify the need for additional risk mitigation requiring verification and/or validation activities?
19 | RISK MANAGEMENT | 3.2 | Engineering Representative | Does this design change introduce new hazards or new causes, and impacts change to the residual risk(s) or their acceptability?
20 | RISK MANAGEMENT | 3.3 | Engineering Representative Risk Management Representative | Does this design change implementation require updates to Risk Assessment and Control Document (RAC), Product Cause and Mitigation Table (CMT)?
21 | RISK MANAGEMENT | 3.4 | Engineering Representative | Does this design change add or change Design Input or other requirements traced to Essential Characteristic or finished Design Output? If Yes, assess impact to product RMF deliverables
22 | RISK MANAGEMENT | 3.5 | Engineering Representative | Does this design change impact software classified as Class B or C? If Yes, access impact to product RMF deliverables
23 | RISK MANAGEMENT | 3.6 | Engineering Representative | Does this design change introduce new product feature or functionality? If Yes access impact to product RMF deliverables
24 | RISK MANAGEMENT | 3.7 | Engineering Representative | Does this design change introduce change to material that is either non-equivalent to the existing material and/or has patient contact (i.e., Biocompatibility)? If Yes, access impact to product RMF deliverables
25 | RISK MANAGEMENT | 3.8 | Engineering Representative | Does this design change add or introduce change to Essential Part (including labeling) and essential controls? If Yes, access impact to product RMF deliverables
26 | RISK MANAGEMENT | 3.9 | Engineering Representative | Does this design change impact product Expected Service Life evaluated and documented in Product Risk Management File/Standard Conformance (Ref: IEC 60601-1 Clause 4.4)?
27 | Failure Modes and Effects Analysis | 4 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to FMEA process? If NO, document the rationale for no impact. Refer to GEHC_GQP_08.01 Risk Management Procedure (DOC0283944) for FMEA Templates & minimum requirements Refer to Failure Modes and Effects Process Standard (DOC2433962) for further guidance
28 | Failure Modes and Effects Analysis | 4.1 | Usability Representative | Does this design change impact released Task Analysis or Use Failure Mode and Effects Analysis (UFMEA) or requires creation of new UFMEA?
29 | Failure Modes and Effects Analysis | 4.2 | Engineering Representative | Does this design change impact released Design FMEA (or other tools/methods, e.g., FTA, Engineering review and/or analysis) or requires developing new DFMEA records?
30 | Failure Modes and Effects Analysis | 4.3 | Manufacturing Representative | Does this design change impact released Process FMEA for manufacturing and/or used medical device processing or requires development of new PFMEA records?
31 | Failure Modes and Effects Analysis | 4.4 | Service Representative Installation Representative | Does this design change impact released Process FMEA for service, installation/de-installation/upgrade or requires development of new PFMEA records for service/installation/maintenance?
32 | Failure Modes and Effects Analysis | 4.5 | Service Representative Installation Representative | Does this design change affect distribution processes requiring special handling (i.e., not able to be transported via general cargo) OR Impacts distribution PFMEA? NOTE: Distribution includes packaging, transportation, and storage.
33 | REGULATORY ASSESSMENT | 5 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to product REGULATORY design controls element? If NO, document the rationale for no impact.
34 | REGULATORY ASSESSMENT | 5.1 | Regulatory Affairs Representative Quality Representative | Does this design change represent Substantial/Significant Change to quality processes or product required per GEHC_GQP_03.01_Quality Plan Procedure (DOC0289782) and warrants communication to Notified Body (NB).
35 | REGULATORY ASSESSMENT | 5.2 | Regulatory Affairs Representative | Does this design change impact Regulatory certifications/registrations/licenses of the end use products? e.g., AERB, CE mark, Health Canada license, China NMPA certs/PTR, country specific license/registrations, Notified body certification (TÜV, BSI, GMED, CSA, UL, etc.)?
36 | REGULATORY ASSESSMENT | 5.3 | Engineering Representative | Does this design change impact Labeling (e.g., updates to system labels, rating plates, regulatory/safety labels, instructions for use, indications for use, addition/deletion of contraindications, GTIN)? NOTE: Change in rating plate may have UDI implications (if it is a new medical device, ensure the medical device is structured at or below the rating plate level in the BOM).
37 | REGULATORY ASSESSMENT | 5.4 | Engineering Representative | Does this design change impact the control mechanism, operating principle, energy type, main electrical ratings, or certification and testing standards?
38 | REGULATORY ASSESSMENT | 5.5 | Engineering Representative | Does this design change impact sterile components, sterile packaging, sterilization methods or changes in biological material of the device?
39 | REGULATORY ASSESSMENT | 5.6 | Manufacturing Representative | Does this design change impact procedures, product manufacturing processes, test specifications/reports, etc. that have regulatory impact?
40 | REGULATORY ASSESSMENT | 5.7 | Engineering Representative | Does this design change impact regional regulatory requirements and safety standards?
41 | REGULATORY ASSESSMENT | 5.8 | Engineering Representative | Does this design change affect MDD/MDR technical file, Declaration of Conformity, 510(k) claims, PMA device, Letter to File, essential principles, annual reports, product report, GSPR, or EU DoC?
42 | REGULATORY ASSESSMENT | 5.9 | Regulatory Affairs Representative | Does this change require the completion of the MDR Significant Change Assessment Form (DOC2193466)? Select appropriate. 1.No.Not Applicable – Product is a Class I product and EU MDR Compliant 2.No: Not Applicable –Product is a Class II product and EU MDR compliant 3.No: Not Applicable – Change does not impact the CE Marked elements of a device (changes to regional device labels outside of the EU – not applicable if changing the device label that bears the CE Mark) 4.No: Not Applicable – Non-significant change based on scope (Changes that do not impact the product such as addition of secondary source supplier, documentation error fixes, component obsolescence change that is not impact form fit and function)" 5.Yes: Attributes as per Chart M of DOC2193466 are impacted. Need evidence of Change assessment form completion either attached to ECO or separate document.
43 | REGULATORY ASSESSMENT | 5.10 | Regulatory Affairs Representative | Does this design change impact established regulatory compliance plan (RCP) or post commercial release RA strategy?
44 | REGULATORY ASSESSMENT | 5.11 | Regulatory Affairs Representative | Does this design change Introduce new commercial catalogs or offerings, requiring RA action?
45 | REGULATORY ASSESSMENT | 5.12 | Regulatory Affairs Representative | Does this design change introduce new accessory that was not previously registered with the device? NOTE: Adding a new medical device accessory according to MDR might trigger RA action.
46 | REGULATORY ASSESSMENT | 5.13 | Regulatory Affairs Representative | Assessment required per DOC0860577 Global RA Change Assessment Guide?
47 | REGULATORY ASSESSMENT | 5.14 | Engineering Representative | Does this design change introduce change to ergonomics, software/firmware, environmental, performance or dimensional specifications? If yes, does this change require the need for clinical data to establish safety and effectiveness or result in design validation that raises new issues of safety and effectiveness?
48 | REGULATORY ASSESSMENT | 5.15 | Engineering Representative Regulatory Affairs Representative | Does this design change include FMI/safety fixes (reportable in any country)? If yes, identify action to perform post market assessment.
49 | REGULATORY ASSESSMENT | 5.16 | Engineering Representative | Other, identify impacted work products and implementation task required.
50 | DESIGN OUTPUT | 6 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN OUTPUT design controls element? If NO, document the rationale for no impact.
51 | DESIGN OUTPUT | 6.1 | Marketing Representative | Does this design change Modify or releases a new commercial offering identified in PCM (Product Configuration Master), ACM (Accessories Configuration Master), ATO (Assemble To Order), PTO (Pick To Order) etc.?
52 | DESIGN OUTPUT | 6.2 | Engineering Representative | Does this design change Impact part(s) that are classified as a GEHC Medical Device? AND the proposed change: Modifies Brand name, Model name OR... modifies UDI published attributes (Product Model number, Manufacturer…) OR... modifies quantity of mandatory component OR... replaces/removes mandatory component OR... changes/removes language from system’s user interface/outputs, labeling, documentation or packaging OR... changes Part Number of Medical Device OR... changes/new medical device identification (certification, registration, regulatory labeling)
53 | DESIGN OUTPUT | 6.3 | Engineering Representative Configuration Management Specialist/Leader | Does this design change Impact any of the Product Controllership attributes including UDI? (including a change to BOM/Collector BOM/Parent Catalog that impacts medical device offerings, or upgrades/replacement parts)
54 | DESIGN OUTPUT | 6.4 | Engineering Representative | Does this change impact I&T (identification and traceability) items? If yes, list impacted I&T part and related I&T Part list document that needs to be revised if applicable, and confirm if this item need to be tracked in global tracking system. (Refer to GEHC_GQP_12.04 Identification and Traceability Procedure (DOC0373356) for further guidance)
55 | DESIGN OUTPUT | 6.5 | Marketing Representative | Does this design change impact Global Data Synchronization Network (GDSN)/GTIN,Commercial Identifiers, Promotional Materials/Brochures, Impact on User Needs and Intended Uses, Product appearance and Product cost? If Yes, confirm with Marketing for any actions to be completed prior to implementation
56 | DESIGN OUTPUT | 6.6 | Engineering Representative Product Security Leader | Does this design change impact Cybersecurity Design Output deliverables: Impacts existing or requires new cybersecurity deliverables (e.g., DEPS Controls Baseline Catalog, Threat Model, SRA, PIA, MMSR, or Security Assessment Summary)? If yes, ensure PSR/PSL approvals are obtained as applicable
57 | DESIGN OUTPUT | 6.7 | Engineering Representative Product Security Leader | Does the design output updates in scope of this design change introduce new cybersecurity vulnerabilities that need to be assessed? Or are there any cybersecurity improvements that need to be implemented with this change?
58 | DESIGN OUTPUT | 6.8 | Engineering Representative | Does this design change impact released Design Specifications (Finished design outputs from design phase) e.g., ADW/DDW/PSP/TST, requiring update?
59 | DESIGN OUTPUT | 6.9 | Engineering Representative | Does this design change impact Bill of Materials for an assembly? If so, the assembly needs to be assessed for opportunities to use standard parts, using the Managed Library Report. See DOC2289045 GEHC Component Selection Guidance
60 | DESIGN OUTPUT | 6.10 | Service Representative | Does this design change impact released product Service Manuals or documentation, e.g., service Tech Pub Labeling? For example, impact to a FRU/DRU part number or service instructions. Refer to Global QS Std WI for Technical Publications (DOC2139328) for further guidance
61 | DESIGN OUTPUT | 6.11 | Installation Representative | Does this design change impact released product Installation manual/Pre-installation/de-installation/upgrade documentation (siting requirements, design center blocks and templates, customer drawings)?
62 | DESIGN OUTPUT | 6.12 | Manufacturing Representative | Is "on-product labeling" printed in-process using a Software program (e.g., BarTender)? If so, does this change impact the following: - Dynamic fields - Static fields - Part verification or manufacturing process validation (e.g., printed label)
63 | DESIGN OUTPUT | 6.13 | Engineering Representative | Does this design change impact released On Product Labeling/Operator/User Manuals (including addendums, datasheet, brochure or any other user facing instructions), e.g., User Tech Pub or Product Labeling/Documentation (Instructions for Use (IFU), translated addendums), product rating plate label, UDI label?
64 | DESIGN OUTPUT | 6.14 | Engineering Representative | Does this design change impact existing Field Replaceable Units (FRU) and/or Depot Replaceable Units (DRU) setup? Does this change require updates in iSpare for example? Is Technical nature attribute set up under classification power view with respect to part nature for Spare Parts?
65 | DESIGN OUTPUT | 6.15 | Engineering Representative Regulatory Affairs Representative | Does this design change Modify or releases new accessory(s)? [Pre-qualified accessories, accessories that can be connected to or interact with the medical device/product may require engineering review to determine need for technical justification or full verification testing or full validation testing, i.e., reference the testing or justification and DHF documentation updates for biocompatibility, safety, EMC, design verification for compatibility, etc.]
66 | DESIGN OUTPUT | 6.16 | Engineering Representative Regulatory Affairs Representative | Is there any change that alters established method of compliance to the basic safety and essential performance requirements in the standard? Part(s) in scope of this change is a critical component? Does this Design Change impact NRTL Listing Report/safety standard test reports, CB Certification and Critical Component List? Example: Impact NMPA registration, PTR (Product Technical Requirement) and PTR mapping to design control and production control
67 | DESIGN OUTPUT | 6.17 | Manufacturing Representative | Does this design change impact a harvested part? Is there an impact to established Harvesting WI, SOP, or process document for the product?
68 | DESIGN OUTPUT | 6.18 | Engineering Representative | Does this design change impact EMI/EMC certifications reports, Suspended Masses, Actual Or Possible Harm To Patient, User, or Service Personnel, etc.?
69 | DESIGN OUTPUT | 6.19 | Engineering Representative | Does this design change impact compliance of WEEE passport(s), RoHS or REACH, Phthalate, DEHP (France), Proposition 65 status, EU_EEA Restriction in use of PFOA, its salts and PFOA-related compounds?
70 | DESIGN OUTPUT | 6.20 | Engineering Representative | Does this design change impact Battery or Battery report? (added/changed/removed/restructured)
71 | DESIGN OUTPUT | 6.21 | Engineering Representative | Does this design change impact Chemical, Physical, Biological characteristics/properties?
72 | DESIGN OUTPUT | 6.22 | Engineering Representative | Does this Design Change impact product appearance? If yes, describe change to product appearance and confirm if review/approval from Industrial design team/representative is required.
73 | DESIGN OUTPUT | 6.23 | Engineering Representative | Is the affected part(s) essential characteristics control plan impacted? OR...Are any of the Parts converted from Non-Essential to Essential classification? OR...Are any design output essential characteristics being changed as a result of the design change? OR...Are NEW parts classified as Essential added to Bill Of Materials (BOM)? NOTE: Document the rationale for changing legacy safety /function/Essential part to Non-Essential.
74 | DESIGN OUTPUT | 6.24 | Engineering Representative | Does this design change impact Telecommunication Registrations and Certifications? (WiFi, UWB, Bluetooth, ZigBee, etc.). Refer to GEHC Engineering Standard for Wireless Product Compliance Planning (DOC2360208) for further guidance.
75 | DESIGN OUTPUT | 6.25 | Engineering Representative | Trade Compliance: Does this design change include a change to the Country Of Origin (COO) for an existing buy part?
76 | DESIGN OUTPUT | 6.26 | Engineering Representative | Is this design change associated with updates to Alternate/Substitue parts? If yes consider all aspects of use (e.g., service, software impact) and identify actions to ensure traceability throughout the entire lifecycle of the product
77 | DESIGN OUTPUT | 6.27 | Engineering Representative | Does this design change impact Seismic certifications or have seismic design impact? If YES, indicate required tasks to complete.
78 | DESIGN VERIFICATION | 7 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VERIFICATION design controls element? If NO, document the rationale for no impact.
79 | DESIGN VERIFICATION | 7.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring verification to confirm that specified requirements are met, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Verification Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with verification results. NOTE: Update to Traceability of Design Input and lower-level requirements to associated Design Verification activities and results may be required.
80 | DESIGN VERIFICATION | 7.2 | Engineering Representative | Does this design change require compatibility testing to confirm forward/backward compatibility of proposed design change with other systems, products, or interfaces?
81 | DESIGN VERIFICATION | 7.3 | Engineering Representative | Does this design change require verification regression testing?
82 | DESIGN VERIFICATION | 7.4 | Engineering Representative | Does this design change require technical analysis/specification analysis (e.g., scenarios where there is no impact to product design, however require engineering analysis to confirm this as part of justification) NOTE: Expert Evaluation and Analysis, including Technical Design Review, FMEA, FTA, and retrospective analysis of historical data applicable to the current design is one of the verification method and technique, per Verification GWI
83 | DESIGN VERIFICATION | 7.5 | Engineering Representative | Does this Design Change make a change to an accessory that would require a Compatibility Assessment with a host/impacted device?
84 | DESIGN VERIFICATION | 7.6 | Engineering Representative | Does this design change impact Engineering test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for engineering development requiring re-verification/revision/re-validation/qualification?
85 | DESIGN VALIDATION | 8 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN VALIDATION design controls element? If NO, document the rationale for no impact.
86 | DESIGN VALIDATION | 8.1 | Engineering Representative | Does this design change impact design inputs (new or existing) requiring validation to confirm that the product continues to meet user needs and intended uses, mitigations are effective, and that no unintended consequences have been introduced as a result of the design change? NOTE: Validation Plan may be identified as a required work product to be completed at ECR phase or identified as a supporting evidence along with validation results. NOTE: Update to Traceability of Design Input requirements to associated Design Validation activities and results may be required. NOTE: Consider all user groups when evaluating impact to validation, i.e., Service, Clinical, etc.
87 | DESIGN VALIDATION | 8.2 | Engineering Representative | Does this design change require validation regression testing?
88 | SOFTWARE | 9 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to SOFTWARE design controls element? If NO, document the rationale for no impact.
89 | SOFTWARE | 9.1 | Engineering Representative | Does this design change impact SOUP/third-party purchased software components? NOTE: If Yes, evaluate impact as per risk management and purchasing controls requirements. NOTE: If Yes, also verify to ensure that SOUP/third-party software functionality and no unintended side effects are introduced.
90 | SOFTWARE | 9.2 | Engineering Representative | Does the review of software (including SOUP) change determines whether: - Additional potential causes are introduced contributing to hazardous situations. - Additional risk mitigations are required. - The software change could interfere with existing risk mitigations. If Yes and there is an impact, identify relevant risk management activities based on impact of software design changes.
91 | SOFTWARE | 9.3 | Engineering Representative | Does this design change require action per released product software maintenance plan (e.g., software architecture, SDD, software release, SOUP Index, software classification)?
92 | SOFTWARE | 9.4 | Engineering Representative | Does this design change include software upgrades, bug fixes, or obsolescence of SOUP items requiring software testing to ensure no unintended side effects are introduced?
93 | SOFTWARE | 9.5 | Engineering Representative | Does this Design Change require released software validation test procedure to be executed to reestablish software validation?
94 | SOFTWARE | 9.6 | Engineering Representative | Does this Design Change require updates to licensing servers/activation code servers based on software versioning?
95 | SOFTWARE | 9.7 | Engineering Representative | Is this software change compatible with all older revision of hardware's in IB? Does this Design Change make a change to the Product Hardware that would require a System Compatibility Assessment?
96 | SOFTWARE | 9.8 | Engineering Representative | Is this software change compatible with all older revision of software in IB? Does this Design Change make a change to the Product Software/Firmware that would require a System Compatibility Assessment/regression test?
97 | USABILITY | 10 | Usability Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to USABILITY design controls element? If NO, document the rationale for no impact.
98 | USABILITY | 10.1 | Usability Representative | Does this design change impact Usability File Index (UFI) deliverables requiring update to UFI or Use Analysis Report (UAR)?
99 | USABILITY | 10.2 | Usability Representative Clinical Representative | Does this design change impact released Device Use Definition (including intended use, patient population, user profile, use environment, and primary operating functions and/or associated Use Scenarios (per UAR))? NOTE: Usability applies to interactions among the patient, User, device, and Use Environment. NOTE: Consider all user groups when evaluating impact to usability, i.e., Service, Clinical, etc.
100 | USABILITY | 10.3 | Usability Representative Clinical Representative | Does this design change impact Use-based hazards/risk mitigations and/or introduces new use-based hazards/risk mitigations?
101 | USABILITY | 10.4 | Usability Representative | Does this design change impact Critical Task or identifies new Critical Task?
102 | USABILITY | 10.5 | Usability Representative | Does this design change impact Formative Evaluations and/or Summative Usability testing? If Yes, Identify action to create Summative Testing Plan and/or Formative Evaluation Plan as applicable.
103 | USABILITY | 10.6 | Usability Representative | Does this design change impact User Interface of Unknown Provenance (UOUP) and is UOUP objective evidence needed (Device Use Definition in Use Analysis Report and documentation of Use Errors identified in Post-Market Data, including, but not limited to complaints, hazardous investigations with root cause of Use Error, near misses, adverse events, and recalls)?
104 | USABILITY | 10.7 | Usability Representative | Does this design change address Use Errors identified in complaints, adverse events, and/or recalls requiring a Known Use Problems evaluation?
105 | DESIGN TRANSFER | 11 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to DESIGN TRANSFER design controls element? If NO, document the rationale for no impact. NOTE: It is recommended to use GEHC_GQP_10.01.007_T002 Design Transfer Plan Template (DOC2317856) for iteration and growth program.
106 | DESIGN TRANSFER | 11.1 | Service Representative | Does this design change impact Service Finished design outputs? (e.g., Service manuals, service notes, asset creation in the service CRM, service checklist, replacement procedures, spare parts list, tools, processes, UDI Label Instructions, technical bulletin) Refer to QS_4004 Global Standard Work Instruction for Item Creation (DOC2473345) for further guidance on asset creation in service CRM.
107 | DESIGN TRANSFER | 11.2 | Service Representative | Does this Design Change require Installation & Service Process Validation? If yes, provide the installation & service process validation plan/procedure document number, and complete validation prior to ECO implementation.
108 | DESIGN TRANSFER | 11.3 | Manufacturing Representative | Does this design change impact MAKE Parts? (i.e., Change to Manufacturing Process/Procedures/Work Instructions(WI)/OMS, SOP or require a New Process/Procedure, Process Validation, in-process label printing, testing, incoming, in-process, final acceptance, tooling, training, e.g., PPC or MVP), MTS, and PCL (Part Classification List)
109 | DESIGN TRANSFER | 11.4 | Environmental Health and Safety (EHS) Representative | Does this design change affect Environmental Health and Safety (EHS) or EHS evaluation?
110 | DESIGN TRANSFER | 11.5 | Supplier Quality Representative | Does this design change require additional external Supplier Qualification or Notification (SCN) for affected parts? NOTE: Check if the ECR introduces parts from an external supplier that is not on the site ASL. NOTE: GEHC Initiated changes may require SCN (Supplier Change Notification) per GEHC_GQP_11.01.009 Purchasing Change Management Work Instruction (DOC0288388) NOTE: Check if SCN needs to be issued to suppliers tasked with procuring the affected parts, if this change is associated with MEP maintenance.
111 | DESIGN TRANSFER | 11.6 | Sourcing Representative Materials Representative | Does this design change impact a part that is EOL (End Of Life) requiring LTB (Last Time Buy) action to be initiated or assessed?
112 | DESIGN TRANSFER | 11.7 | Manufacturing Representative | Does this design change require Incoming Material Qualification for BUY Parts (Manufacturing and Drop Ship)? Does Incoming acceptance document (e.g., ITD), Characteristic Card(s), PID need to be created/modified/removed? (required for part with purchasing or incoming manufacturing controls)?
113 | DESIGN TRANSFER | 11.8 | Engineering Representative | Does this design change impact established Product Monitoring design transfer strategy: Requires update to Product complaint handling process/system (consider impact to PSI codes and associated information, e.g., PSI Code, PSI Description, LEGAL MFG, MFG SITE, DESIGN OWNER)? OR...Requires activities to define content of reporting e.g. regulatory traceable device identification, Device History Record and Inventory Control, and Annual Report Identification and Set-Up? OR...Requires activities to define GEHC site interactions, e.g., update Site ASL with new GEHC Site, Quality Agreement, oversight and monitoring methodologies?
114 | DESIGN TRANSFER | 11.9 | Manufacturing Representative | Does this design change affect eDHR or Paper DHR?
115 | DESIGN TRANSFER | 11.10 | Engineering Representative | Does this design change introduce a new or update to manufacturing Tool/software?
116 | DESIGN TRANSFER | 11.11 | Engineering Representative | Does this design change impact released software delivery process? (e.g., physical media, duplication process, eDelivery process) NOTE: Ensure the released software can be reliably delivered from its archive to the point of use without corruption or unauthorized change.
117 | DESIGN TRANSFER | 11.12 | Supplier Quality Representative | Do any of the impacted parts require Purchasing controls for established Essential Characteristics?
118 | DESIGN TRANSFER | 11.13 | Supplier Quality Representative | Does PQP need to be updated for any other reason besides Essential controls?
119 | DESIGN TRANSFER | 11.14 | Supplier Quality Representative | Does this design change require new jigs/tooling/mould at the supplier, or is any of the existing jigs/tooling/mould at the supplier affected by this change?
120 | DESIGN TRANSFER | 11.15 | Supplier Quality Representative | Does this design change impact Supplier test equipment, tools, simulators, software, product bays, products (used for testing/equipment qualification, etc.) used for manufacturing testing requiring update? If Yes, SCR is Required from Supplier for test equipment
121 | DESIGN TRANSFER | 11.16 | Manufacturing Representative | Does this design change impact Production Process Validation, Equipment qualification, Tooling or Fixture verification, Manufacturing Software, Training, Packaging and Shipping (e.g. labeling, shelf life, storage requirements). Impacts Material handling procedures, Environmental Requirements (Shelf Life, Temperature, Humidity Controls, etc.) New chemical introduced/SDS impacted, Manufacturing Process Capability
122 | DESIGN TRANSFER | 11.17 | Supplier Quality Representative | Is the critical supplier not identified in current critical supplier list? If Yes, identify action for Engineering to update critical supplier list as needed.
123 | DESIGN TRANSFER | 11.18 | Supplier Quality Representative | Is the critical supplier list not updated? If Yes, identify action for SQE to update ASL and notify suppliers of the critical supplier designation.
124 | DESIGN TRANSFER | 11.19 | Supplier Quality Representative Sourcing Representative | Is the critical supplier list not updated? If Yes, identify action for Sourcing to update SQR (Supplier Quality Requirements).
125 | DESIGN TRANSFER | 11.20 | Engineering Representative | Is the critical supplier list not updated? If Yes, identify action for RA to review/update Product Technical File as needed.
126 | DESIGN TRANSFER | 11.21 | Manufacturing Representative | Does this design change impact Dedicated Test Equipment (including SW options media stations)?
127 | DESIGN TRANSFER | 11.22 | Materials Representative | Is a buy-level part supplied by another GEHC site introduced with this change? NOTE: Check if the ECR introduces parts from a GEHC affiliate/supplier that is not on the site ASL. NOTE: Consider any updates required to the Site Product Matrix for new GEHC suppliers or new services provided by GEHC suppliers
128 | DESIGN TRANSFER | 11.23 | Service Representative | Does this design change impact Proprietary Classification of Hardware/Software Tools, Documentation, and Files documented in Product Proprietary Classification record per QS_10001 Global Standard WI for Proprietary Classification (DOC1410674)?
129 | DESIGN TRANSFER | 11.24 | Materials Representative | Does this design change impact Buy Parts? (Change To/Switching Supplier, New Direct Supplier Identified (Confirm Addition in ASL),New Sub-Tier Supplier Identified Part cost changing , Supplier Tooling changing) Does part categorization match supplier categorization, Approved Supplier List, Blanket Purchase Agreement, Purchase Order?
130 | Cleaning, Disinfection, and Sterilization | 12 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to Cleaning, Disinfection, and Sterilization design controls element? If NO, document the rationale for no impact.
131 | Cleaning, Disinfection, and Sterilization | 12.1 | Engineering Representative | Does this design change impact chemicals listed in the user manual? Check https://cleaning.gehealthcare.com/ to ensure the chemicals listed match the GE Healthcare cleaning website.
132 | Cleaning, Disinfection, and Sterilization | 12.2 | Engineering Representative | Does this design change impact established product cleaning and disinfection processes? If the design, materials, or device use has changed substantially, then a new validation or adoption assessment may need to be performed. Use GEHC_GQP_10.01.019_T001 Cleaning and Disinfection Plan Template (DOC2440136) [Device use change - if the device use has changed, then the Spaulding classification needs to be performed. If the Spaulding classification changes, then the cleaning and disinfection processes may need to be changed. Material change - probably does not affect the cleanability or disinfection but may have an effect on the material compatibility with the chemical. Design feature change - the change needs to be compared to the previous design feature to see if this is more difficult to clean and disinfect.] NOTE: Any addition or removal of cleaning agents requires assessment of Design input(SRS), Design output(Manuals), and Website.
133 | TECHNICAL DESIGN REVIEW | 13 | Engineering Representative | Do any of the IMPACT ANALYSIS considerations/questions below apply to TECHNICAL DESIGN REVIEW design controls element? If NO, document the rationale for no impact. Refer to QS_4002 Global Standard WI for Technical Design Review (DOC0707915) for further guidance.
134 | TECHNICAL DESIGN REVIEW | 13.1 | Engineering Representative | Does this proposed design change implementation require TDR to be planned and reviewed by engineering SME (e.g., Chief Engineer, LSD, PEs, Architects)? If yes identify TDR as part of TDR Calendar or implementation action plan and identify required SME based on topic. Examples may include Reliability TDR, Safety TDR, Usability TDR, Manufacturability TDR, Serviceability TDR, Software TDR, Compatibility TDR, Interoperability TDR.
135 | TECHNICAL DESIGN REVIEW | 13.2 | Change Owner or Design Owner | Does this design change require TDR to be planned and executed per QS_3001 Global Standard WI for Suspended Mass Design (DOC1521734)?
136 | TECHNICAL DESIGN REVIEW | 13.3 | Engineering Representative Product Security Leader | Does this design change require TDR to be planned and executed per GEHC New Product Cybersecurity Standard (DOC2102581)?
137 | TECHNICAL DESIGN REVIEW | 13.4 | Engineering Representative | Does this design change require TDR to be planned and executed per GEHC New Product AI Quality Standard (DOC2252997)?
138 | TECHNICAL DESIGN REVIEW | 13.5 | Service Representative | Does this design change require TDR to be planned and executed per QS_1001 Global QS Standard WI for Product Service Standards (DOC0487759)? If YES Service TDR Guidance and Checklist (DOC2444001) should be used.
139 | TECHNICAL DESIGN REVIEW | 13.6 | Engineering Representative | Does this design change require TDR to be planned and executed per GEHC Design Standard For Rechargeable DC Energy Storage Devices (Secondary Batteries and Super Capacitors) (DOC2385440)? NOTE: Design change impacting Safety and Charging parameters, storage and shelf life, portable power management, Portable power capacity, updates to PSP etc., may require TDR to be planned and approved before implementing design change.
140 | TECHNICAL DESIGN REVIEW | 13.7 | Engineering Representative | Are any additional reviews required? - Purchased GE Designed Parts/Assemblies - Design Requirements Review with Supplier - Documented Cross Functional Core Team Review - Manufacturing Review, - Regulatory Review - Marketing/Product Management or Clinical Review - Intellectual Property Review - Require Drawing review if this design is from Predicate device ? - Require Tolerance analysis for critical joints/assemblies ?
141 | PART(S) DISPOSITION | 14 | Materials Representative | Actions to disposition impacted materials and product on-order, in-process, manufacturing, distribution, service spares (e.g., use-up, rework, scrap, no change, balance) is required?
142 | PART(S) DISPOSITION | 14.1 | Materials Representative | Does this Design Change Plan identify parts with "rework" and "scrap" disposition? If Yes, include action to identity rework instructions and capture evidence of rework and scrap as part of ERP Release or ECO implementation
143 | PART(S) DISPOSITION | 14.2 | Engineering Representative | Does this design change require Hold, Concession or Deviation to be initiated? If Yes, identify traceability to document number of Hold, Concession or Deviation.
144 | PART(S) DISPOSITION | 14.3 | Marketing Representative | Does this design change obsolete/end of life parts or catalogs or software or configurations or offerings as part of Product M6-M9 PRD milestone phase? If Yes, identify tasks to ensure impacts items are also obsoleted from order management/Marketing systems (e.g., Oracle, Salesforce).
145 | Other RDOs | 15 | Engineering Representative | Does this design change impact another modality, system or subsystem managed by another change board If NO, document the rationale for no impact. If YES, indicate required tasks to complete.
146 | Platform Software or Hardware | 16 | Engineering Representative | Does this design change impact Platform software or hardware? If yes, identify design controls deliverables required to be completed prior to change implementation (e.g., platform integration and integration testing activities) If NO, document the rationale for no impact.
147 | OTHERS